## VB12
Sharada Sivaraman
2012-06-07
* In bold are time estimates for future experiments
Vitamin B12
| Experiment | Status | Timeline* | Details
| Resource |  | Days
1 | MTP screen |  | 3 months | Expt 1 and 6 are done side by side to evaluate which assay format and binder combination will yield the best results. At first only AB calibrators were used
more recently 4 clinicals are being tested along with sample treatment/extraction procedures to finalize format. Please see issue #1 and #2
Two commercial anti B12 antibodies are also being evaluated (i) Millipore and (ii) from Monobind
2 | Conjugation |  | 10 days | Six sources of I.F protein were conjugated with BT, 2 of them were also conjugated with AP. 3 different anti-IF clones were BT and AP conjugated..
The recovery for the IF protein conjugations was poor on Dojindo kit and the Invitrogen kit was used
3 | Cross reactivity |  | 5 | Six sources of I.F protein were conjugated with BT, 2 of them were also conjugated with AP. 3 different anti-IF clones were BT and AP conjugated..
4 | Interference |  | 5 | Same as above
5 | Epitope map | N/A
6 | Theranos screen |  | 2 months | Same as experiment 1. The clinical samples are in serum matrix. After finalization of binder/format will need to run std curves in normal plasma , normal serum, depleted serum. See issue #4
7 | Training set |  | 7 days | Around N=70 clinical samples and B12 calibrators have been evalaued on the Siemens Immulite VB12 assay. Training set pending finalization of assay format/binder
| Antibody Pair finalized |  | 2 months | This will take additional 2 months provided protocols are available this week
8 | Plasma Screen |  | 5
9 | Whole Blood Screen | N/A
10 | Capture titration |  | 5
11 | Detection titration |  | 2
12 | Detection Stabilizer |  | 2
13 | Effect of Buffer |  | 3 weeks | Sample treatment uses 3 buffers: High pH buffer containing KCN and cobinamide, DTT and neutralization buffer. Each buffers' composition and pH will need to be optimized.
Assay diluent and coating buffer will also need optimization. Stabilizers/additives will need to be evaluated,.
14 | Effect of Incubation time |  | 3 weeks | This will include additional protocol requests modifying the B12 protocols further for reagent incubation times.
15 | Reader protocol optimization |  | 3 weeks | Please see issue #3. In the interim off board sample treatment is being done.
16 | Calibrator comparison |  | 5
17 | Whole blood spike recovery |  | 5
18 | Hematocrit effect
19 | Plasma spike recovery
20 | LLOQ/ULOQ |  | 5
21 | Matrix effects |  | 5 | B12 spiked into interfering matrices will need to be validateds on Siemens Immulite assay as well as Theranos assay
22 | Specificity: RF, HAMA |  | 2 | same as above
23 | Clinical samples |  | 10
24 | Stability |  | 12 weeks | After expt 13, stability of reagents will need to be carried out and troubleshooting will need ot be done on a reagent by reagent basis
| Issues | 1 | Despite screening 14 different assay conditions that include various binder combinations (including binding protein (intrinsic factor) from 6 sources (both native and recombinant), B12BSA conjugate, Anti-IF antibody)
low modulations of S/B 2 are routinely seen. A checkerboard titration of all reagents (binder and AP conjugate) needs to be performed.
|  | 2 | Out of the 6 sources of the binding protein it was discovered that 3 already contain B12 as a result of production at the vendors end. These are the recombinant versions and they have since been excluded form screening.
|  | 3 | The B12 assay sample treatment protocol requires 3 steps. It is highly desirable to have a protocol that would do this on board. A request for 4 protocols was made on 5/22/12. No protocols have been received to date.
|  | 4 | B12 levels tend to be high to mid range in normal donors. It will be hard to find matrix with low levels to make spiekd calibrators. Depleted B12 sera also not available. Folate depleted serum has 240 pg/mL of B12
, we need to target an LLLOQ of 100 pg/ml.
| Completed In |  | 4 months from today | If binder and assay format picked show acceptable correlation with Siemens Immulite for clinical samples

## Methadone
Sharada Sivaraman
2012-06-07
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen
2 | Conjugation
3 | Cross reactivity
4 | Interference
5 | Epitope map
6 | Theranos screen
7 | Training set
Antibody Pair finalized
8 | Plasma Screen
9 | Whole Blood Screen
10 | Capture titration
11 | Detection titration
12 | Detection Stabilizer
13 | Effect of Buffer
14 | Effect of Incubation time
15 | Reader protocol optimization
16 | Calibrator comparison
17 | Whole blood spike recovery
18 | Hematocrit effect
19 | Plasma spike recovery
20 | LLOQ/ULOQ
21 | Matrix effects
22 | Specificity: RF, HAMA
23 | Clinical samples
24 | Stability
| Issues |  | 1 | More antibodies and reagents being ordered.
|  |  | 2 | Awaiting Methadone-AP conjugates from synthesis group
|  |  | 3 | Will be able to estimate timelines better after receipt of reagents
Completed In

## HbsAg
HBsAg Assay
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen |  | 3 weeks | Preliminary screening of sandwich pairs on MTP, 7x7x4, 21 anitbodies pairs.
2 | Conjugation |  | 2 | Biotin and AP conjugation of all Antibodies
3 | Cross reactivity |  | 3 | Screen antigens that include: HIV, HCV, CMV, HAV, EBV etc.
4 | Interference |  | 2
5 | Epitope map |  | 2
6 | Theranos screen |  | 1 week | Tip coating request for above capture combinations. Tips dried O/N will be used to screen on Reader with Wildtype and mutants antigens.
7 | Training set |  | 1 week | Test antigen positive samples in BSL2 lab with the available Ab combinations so far . Evaluate all Ab combinations that are promising so that backup can also be finalized.
Antibody Pair finalized
8 | Plasma Screen |  | 1 week
9 | Whole Blood Screen |  | 1 week
10 | Capture titration |  | 2 | Assay optimization
11 | Detection titration |  | 2 | Assay optimization
12 | Detection Stabilizer |  | 2 | Assay optimization
13 | Effect of Buffer |  | 2 | Assay optimization
14 | Effect of Incubation time |  | 2 | Assay optimization
15 | Reader protocol optimization |  | 2 | Assay optimization
16 | Calibrator comparison |  | 2
17 | Whole blood spike recovery |  | 2
18 | Hematocrit effect |  | 2 | Screen whole blood samples for samples with low levels of HbSAg. Whole blood spike recovery, plasma from spiked whole blood: Hematocrit
19 | Plasma spike recovery |  | 2
20 | LLOQ/ULOQ |  | 2
21 | Matrix effects |  | 3 | Spike normal serum, plasma and whole blood with WT and mutant antigen material (that has been verified on Siemens Immulite)
22 | Specificity: RF, HAMA |  | 3
23 | Clinical samples |  | 4 weeks | 7 panels positive samples
24 | Stability |  | 12 weeks | stability for tips and Detections
Issues
| Completed In |  |  | Screening both combination 2 and combination 5 and also binders with all above conditions.

## Folate
Folic acid assay development
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen
2 | Conjugation |  |  | A few conjugates being used in early development stage were specially designed and produced by in-house chemistry group. Materials selection was done but new batches of conjugates will be evaluated in later stage.
3 | Cross reactivity |  | 3
4 | Interference |  | 3
5 | Epitope map |  | N/A
6 | Theranos screen |  | 3
7 | Training set |  |  | Some clinical samples were used for sample treatment procedure development and assay format evaluation.
Antibody Pair finalized
8 | Plasma Screen |  | 3
9 | Whole Blood Screen |  | N/A
10 | Capture titration |  | 2
11 | Detection titration |  | 2
12 | Detection Stabilizer |  | 3
13 | Effect of Buffer |  |  | This assay requires a sample pretreatment. Huge effort has been put into optimize sample treatment combining with assay format design/test.
14 | Effect of Incubation time |  | 3 | This is on-going together with protocol optimization.
15 | Reader protocol optimization |  | 3 | This assay requires sample pretreatment procedure. New Edison protocols were obtained and are still under optimization.
16 | Calibrator comparison |  | 3 | The requirement of this assay is to measure both Folic Acid and 5mTHF, so calibrators have to evaluated for both Folic Acid and 5mTHF measurement.
17 | Whole blood spike recovery |  | N/A
18 | Hematocrit effect |  | N/A
19 | Plasma spike recovery |  | 2
20 | LLOQ/ULOQ |  | 2
21 | Matrix effects |  | 3
22 | Specificity: RF, HAMA |  | 2
23 | Clinical samples |  | 3
24 | Stability |  | 3 months
Issues
1. A sample treatment procedure is critical to this assay development.
2. Assay development was transited from MTP to Edison early due to sample treatment method development.
| Completed In |  | 2 months except stability

## Transferrin
Assay | Transferrin (TRF)
Developer | Amy Dinh
eLog | 0722
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen |  |  | 20x20 Ab pairs were screened on MTP, initially, 4 best pairs selected.
2 | Conjugation |  |  | Biotin-SH labeled, then AP-SH labeled for all 20 Abs. Note: Conjugation occurs every so often to replenish conjugate inventory.
3 | Cross reactivity |  |  | Started on MTP, but observed saturation, so switched to Theranos System. Additional pairs selected for assay specificity and awaiting Abs arrival.
However,one Ab pair lost S/B modulation, and had to step-back to trouble-shoot.
4 | Interference |  |  | Conjugated more for tip coating and Tray QC'ing. Additional pairs selected for assay specificity and awaiting Asb arrival.
5 | Epitope map | N/A
6 | Theranos screen |  |  | Current protocol: Transferrin_10000x_1-1-1 svn 6742. Matrix chosen: assay buffer.
7 | Training set |  |  | Ab pairs predicted sample's conc only 1/2 what the SIEMENS_ADVIA predicted. So need to calculate the % Difference from Target.
| Antibody Pair finalized |  | 7 | TBD
8 | Plasma Screen |  | 1-2
9 | Whole Blood Screen |  | 1-2
10 | Capture titration |  | 1-2
11 | Detection titration |  | 1-2
12 | Detection Stabilizer |  | 1-2
13 | Effect of Buffer |  | 1-2
14 | Effect of Incubation time |  | 3-7 | Current protocols: Transferrin_10000x_1-1-1 svn 6742 and Generic2_1000x svn 5742 (this is used for the 3 recently selected pairs).
15 | Reader protocol optimization |  | 3-5 | Lead time is required when requesting new protocols and/or requesting "minor changes" to existing protocols.
16 | Calibrator comparison |  | 2-3 | Need to prepare calibrators, then test them before release for use in future expts.
17 | Whole blood spike recovery |  | 2-3
18 | Hematocrit effect |  | 1-2
19 | Plasma spike recovery |  | 1-2
20 | LLOQ/ULOQ |  | 2-3
21 | Matrix effects |  | 2-3
22 | Specificity: RF, HAMA |  | 1-2
23 | Clinical samples |  | 3-5 | 125+ clinical samples sent to in-house CLIA lab for analysis (SIEMENS_ADVIA) and most fell into the narrow reference range of 200-400 mg/dL.
24 | Stability |  | spans over 12 weeks | Studies can only be started once optimal conditions of assay finalized.
| Issues |  |  | 1. Observed many hits, need to select pairs that are TRF assay specific (i.e. do not cross react with potential similar proteins).
2. Additional 3 pairs selected for testing assay specificity; awaiting Ab arrivals (POs submitted and approved 6-7-12).
3. Need to conduct cross reactivity and interference expts for new pairs.
Please note: this is the best estimate of the # of days required to carry-out the said expt successfully. From time to time, we do experience "issues" that
arise unexpectedly i.e. low in stock reagents and/or supplies (readers, dry generics, conjugation kits, etc). Or there will be times when more urgent and
higher priority special projects that may arise.
| Completed In |  |  | Currently projected completion date of August 2012.

## Scl-70
Assay: | Anti Scl-70
Developer: | Christopher Loi
Date: | 2012-05-07
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen | N/A
2 | Conjugation |  | 7 | Scl70 antigen conjugated using three different BT labeling kits, only the Invitrogen kit provided atleast a 60% yield. Capture surface optimization: Tips coated with BT conjugate and direct coat. Unlabeled antigen proved to be the better option as capture for Scl70 antibody assay.
3 | Cross reactivity |  | 2 | 11 CDC ref sera to be tested
4 | Interference |  | 2 | Screened RF & HAMA on INOVA kit for interferences. Results negative. Pending testing on the Theranos assay
5 | Epitope map | N/A
6 | Theranos screen |  | 2 | Started screening N=15 Normal sera samples but high %CVs possibly attributed to tip coating/formulation-related issue is currently under investigation.
7 | Training set |  | 10 | N=72 samples screened so far on 3 benchtop ELISAs: INOVA, Immco and IBL international
| Antibody Pair finalized |  | 2 | Same as Experiment 2
8 | Plasma Screen |  | 5 | Completed N=15 Normals screening on 3 ELISA kits; Pending testing on Theranos assay
9 | Whole Blood Screen | N/A
10 | Capture titration |  | 2 | Unlabeled Direct Coat provided the best modulation
11 | Detection titration |  | 2 | 100 ng/mL Dab provided the best modulation
12 | Detection Stabilizer |  | 2 | 100 ng/mL Dab in Biostab stabilizer provided the best modulation
13 | Effect of Buffer |  | 2 | Blocking Buffer (without HBR) provided the best modulation
14 | Effect of Incubation time |  | 2 | 10X_PSW provided the best modulation; However, may need to revisit this since this test was done with tips using formulation that failed Tray QC
15 | Reader protocol optimization |  | 2
16 | Calibrator comparison |  | 2
17 | Whole blood spike recovery | N/A
18 | Hematocrit effect | N/A
19 | Plasma spike recovery | N/A
20 | LLOQ/ULOQ |  | 2 | This being a qualitative assay will include computing a cut off. Same as normal plasma screen
21 | Matrix effects |  | 2 | HAMA/RF pos: N=6 samples of each type screened on INOVA kit; Pending testing on Theranos assays
22 | Specificity: |  | 2 | Pending testing of 11 CDC ref sera on Theranos readers
23 | Clinical samples |  | 2 | A final set of N=30 samples to be tested on Theranos assay
24 | Stability |  | 12 weeks | Will be kickstarted as soon as validation expts 20-23 provide satisfactory results
| Issues |  |  | Currently investigating situation with tip coating which has led to high %CVs. Will need to coat new tips and possibly may need to revisit the following: (1) Effect of Incubation Time (2) Effect of Sample Dilution and (3) Screening of Normal Stanford Sera (n=15)
| Completed In |  | 4 weeks ( week of 7/2) | This will include completion of all experiments, start of stability and completion of all reports. Longer time is needed if specificity and clinic al correlation are off.

## D-dimer
Xiaoyan works on D-dimer assay
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen |  | 2 or 3 weeks | Screening on 18 pairs of anti-d-dimer antibodies. All the antibodies we have recognize both d-dimer and d-monomer. Wait for antibodies more specific for D-dimer .
2 | Conjugation |  |  | Both biotin and AP conjugations
3 | Cross reactivity |  | 2 days
4 | Interference |  | 2 days
5 | Epitope map | N/A
6 | Theranos screen
7 | Training set
Antibody Pair finalized
8 | Plasma Screen
9 | Whole Blood Screen
10 | Capture titration
11 | Detection titration
12 | Detection Stabilizer
13 | Effect of Buffer
14 | Effect of Incubation time
15 | Reader protocol optimization
16 | Calibrator comparison
17 | Whole blood spike recovery
18 | Hematocrit effect
19 | Plasma spike recovery
20 | LLOQ/ULOQ
21 | Matrix effects
22 | Specificity: RF, HAMA
23 | Clinical samples
24 | Stability
| Issues |  |  | Whether specific antibodies can be found timely is deciding the assay progess.
Completed In

## Influenza A
Sheena M
Influenza A
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen |  | 10 | Screening the 15 different Captures with each of the 15 different Detection Antibodies. Still need remaining antibody shipment to arrive.
2 | Conjugation |  | 5 | Need to conjugate all 30 antibodies. Still waiting on the remaining antibodies to test
3 | Cross reactivity |  | 10 | Test cross reactivity on MTP-Waiting on items to arrive from NIBSC (United Kingdom)-International shipment
4 | Interference |  | 5 | Test intereference on MTP Waiting on items to arrive from NIBSC (United Kingdom)--International shipment
5 | Epitope map |  | N/A
6 | Theranos screen |  | 5 | Tip coating and Testing on edison with some controls
7 | Training set |  | 7 | Testing clinical samples on the edison and possibly troubleshooting
| Antibody Pair finalized |  | 10 | Bulk order for final pairs
8 | Plasma Screen |  | N/A
9 | Whole Blood Screen |  | N/A
10 | Capture titration |  | 3 | Testing for optimal capture antibody concentration
11 | Detection titration |  | 3 | Testing different detection concentrations
12 | Detection Stabilizer |  | 2 | Testing different detection stabilizers
13 | Effect of Buffer |  | 10 | Development of extraction buffer. Will test our current protocols or I will have to develop a new one
14 | Effect of Incubation time |  | 2 | Testing the different incubation protocols
15 | Reader protocol optimization |  | 5 | Might need a new protocol for this format based on extraction buffer development
16 | Calibrator comparison |  | N/A
17 | Whole blood spike recovery |  | N/A
18 | Hematocrit effect |  | N/A
19 | Plasma spike recovery |  | N/A
20 | LLOQ/ULOQ |  | 2 | Determine our sensitivity range for comparison to other FDA approved kits
21 | Matrix effects |  | 5 | Run on Edison and possibly test with commercial kits
22 | Specificity: RF, HAMA |  | 5 | Run on Edison and possibly test with commercial kits
23 | Clinical samples |  | 10 | In House studies possibly to obtain negative swabs. Compare with commerical vs Theranos system. Final validation study
24 | Stability
Issues
Still waiting on antibodies for initial MTP Screen
Bulk order and shipment for finalized Antibody pairs
Pair should be good and should correlate with clinicals- Start back at Step 1
International shipment airport delays for key reagents
Extraction buffer optimizations with Edison protocol
Cross reactivity issues-Start back at Step 1
Completed In

## Influenza B
Sheena M
Influenza B
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen |  | 10 | Screening the 15 different Captures with each of the 15 different Detection Antibodies. Still need remaining antibody shipment to arrive.
2 | Conjugation |  | 5 | Need to conjugate all 30 antibodies. Still waiting on the remaining antibodies to test
3 | Cross reactivity |  | 10 | Test cross reactivity on MTP-Waiting on items to arrive from NIBSC (United Kingdom)-International shipment
4 | Interference |  | 5 | Test intereference on MTP Waiting on items to arrive from NIBSC (United Kingdom)--International shipment
5 | Epitope map |  | N/A
6 | Theranos screen |  | 5 | Tip coating and Testing on edison with some controls
7 | Training set |  | 7 | Testing clinical samples on the edison and possibly troubleshooting
| Antibody Pair finalized |  | 10 | Bulk order for final pairs
8 | Plasma Screen |  | N/A
9 | Whole Blood Screen |  | N/A
10 | Capture titration |  | 3 | Testing for optimal capture antibody concentration
11 | Detection titration |  | 3 | Testing different detection concentrations
12 | Detection Stabilizer |  | 2 | Testing different detection stabilizers
13 | Effect of Buffer |  | 10 | Development of extraction buffer. Will test our current protocols or I will have to develop a new one
14 | Effect of Incubation time |  | 2 | Testing the different incubation protocols
15 | Reader protocol optimization |  | 5 | Might need a new protocol for this format based on extraction buffer development
16 | Calibrator comparison |  | N/A
17 | Whole blood spike recovery |  | N/A
18 | Hematocrit effect |  | N/A
19 | Plasma spike recovery |  | N/A
20 | LLOQ/ULOQ |  | 2 | Determine our sensitivity range for comparison to other FDA approved kits
21 | Matrix effects |  | 5 | Run on Edison and possibly test with commercial kits
22 | Specificity: RF, HAMA |  | 5 | Run on Edison and possibly test with commercial kits
23 | Clinical samples |  | 10 | In House studies possibly to obtain negative swabs. Compare with commerical vs Theranos system. Final validation study
24 | Stability
Issues
Still waiting on antibodies for initial MTP Screen-Problem currently on vendor side
Bulk order and shipment for finalized Antibody pairs
Pair should be good and should correlate with clinicals- Start back at Step 1
International shipment airport delays for key reagents
Extraction buffer optimizations with Edison protocol
Cross reactivity issues-Start back at Step 1
Completed In

## AFP
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen | started | 8 | Find best antibody pairs for assay.
2 | Conjugation | started | 3 | Label Abs with AP and Biotin.
3 | Cross reactivity | not started | 4 | Make sure Abs don't cross react.
4 | Interference | not started | 4 | Test interfering substances.
5 | Epitope map | not started | 3 | Determine binding site of Abs.
6 | Theranos screen | not started | 3 | Coat tips with best Abs.
7 | Training set | not started | 3 | Test Abs with clinical samples.
| Antibody Pair finalized | not started
8 | Plasma Screen | not started | 3 | Screen plasma samples.
9 | Whole Blood Screen | not started | 3 | Screen whole blood samples.
10 | Capture titration | not started | 3 | Optimize capture concentration.
11 | Detection titration | not started | 3 | Optimize detector concentration.
12 | Detection Stabilizer | not started | 3 | Optimize detection stabilizer.
13 | Effect of Buffer | not started | 3 | Check coating buffers.
14 | Effect of Incubation time | not started | 3 | Optimize assay incubation time.
15 | Reader protocol optimization | not started | 3 | Optimize reader protocols.
16 | Calibrator comparison | not started | 3 | Test other calibrators.
17 | Whole blood spike recovery | not started | 3 | Test recovery in whole blood.
18 | Hematocrit effect | not started | 3 | Test for hematocrit effect.
19 | Plasma spike recovery | not started | 3 | Test recovery in plasma.
20 | LLOQ/ULOQ | not started | 3 | Determine LLOQ/ULOQ.
21 | Matrix effects | not started | 3 | Test for matrix effects.
22 | Specificity: RF, HAMA | not started | 3 | Test RF and HAMA samples.
23 | Clinical samples | not started | 3 | Determine clinical correlation.
24 | Stability | not started |  | Test stability of tips and Dabs.
Issues
The AFP assay will be quantitative and the goal is to measure AFP in serum and amniotic fluid.
Completed In

## Urine hCG
Urine hCG Assay
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen
2 | Conjugation
3 | Cross reactivity
4 | Interference
5 | Epitope map
6 | Theranos screen
7 | Training set |  | 5
Antibody Pair finalized
8 | Plasma Screen
9 | Whole Blood Screen
10 | Capture titration
11 | Detection titration
12 | Detection Stabilizer
13 | Effect of Buffer
14 | Effect of Incubation time
15 | Reader protocol optimization
16 | Calibrator comparison
17 | Whole blood spike recovery
18 | Hematocrit effect
19 | Plasma spike recovery
20 | LLOQ/ULOQ
21 | Matrix effects
22 | Specificity: RF, HAMA
23 | Clinical samples
24 | Stability
Issues
Completed In

## RF IgM
Urine hCG Assay
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen
2 | Conjugation
3 | Cross reactivity
4 | Interference
5 | Epitope map
6 | Theranos screen
7 | Training set |  | 5
Antibody Pair finalized
8 | Plasma Screen
9 | Whole Blood Screen
10 | Capture titration
11 | Detection titration
12 | Detection Stabilizer
13 | Effect of Buffer
14 | Effect of Incubation time
15 | Reader protocol optimization
16 | Calibrator comparison
17 | Whole blood spike recovery
18 | Hematocrit effect
19 | Plasma spike recovery
20 | LLOQ/ULOQ
21 | Matrix effects
22 | Specificity: RF, HAMA
23 | Clinical samples
24 | Stability
Issues
Completed In

## Jo-1
Nahal G. | (Training with SS) |  |  |  | Nahal G. | (New Assay)
| Anti SSA Completed Assay |  |  |  |  | ENA (Jo-1) Assay Development
| Experiment | Status | Timeline | Details |  | Experiment | Status | Timeline | Details
| Resource |  | Days |  |  | Resource |  | Days
1 | MTP screen |  |  |  | 1 | MTP screen
2 | Conjugation |  | 7 |  | 2 | Conjugation |  | 7 | AP conjugated, different BT labeling kits to test
Capture surface optimization : Tips coated with BT conjugate and direct coat. Unlabeled antigen proved to be the better option as capture for Scl70 antibody assay
3 | Cross reactivity |  |  |  | 3 | Cross reactivity
4 | Interference |  |  |  | 4 | Interference
5 | Epitope map |  |  |  | 5 | Epitope map | N/A
6 | Theranos screen |  |  |  | 6 | Theranos screen |  | 3 | Prep and aliquot out samples and start screening n=20 Normal sera samples.
7 | Training set |  |  |  | 7 | Training set
| Antibody Pair finalized |  |  |  |  | Antibody Pair finalized
8 | Plasma Screen |  |  |  | 8 | Plasma Screen
9 | Whole Blood Screen |  |  |  | 9 | Whole Blood Screen
10 | Capture titration |  |  |  | 10 | Capture titration
11 | Detection titration |  |  |  | 11 | Detection titration
12 | Detection Stabilizer |  |  |  | 12 | Detection Stabilizer
13 | Effect of Buffer |  |  |  | 13 | Effect of Buffer
14 | Effect of Incubation time |  |  |  | 14 | Effect of Incubation time
15 | Reader protocol optimization |  |  |  | 15 | Reader protocol optimization
16 | Calibrator comparison |  |  |  | 16 | Calibrator comparison
17 | Whole blood spike recovery |  |  |  | 17 | Whole blood spike recovery
18 | Hematocrit effect |  |  |  | 18 | Hematocrit effect
19 | Plasma spike recovery |  |  |  | 19 | Plasma spike recovery
20 | LLOQ/ULOQ |  |  |  | 20 | LLOQ/ULOQ
21 | Matrix effects |  |  |  | 21 | Matrix effects
22 | Specificity: RF, HAMA |  |  |  | 22 | Specificity: RF, HAMA
23 | Clinical samples |  |  |  | 23 | Clinical samples |  | 3 | A set of N=14 samples have arrived (Bioreclamation and Promeddx) but more may need to be order?
24 | Stability |  |  |  | 24 | Stability
| Issues |  |  |  |  | Issues (New Assay):
| Training with SS |  | 3 | testing of pos and neg ctrls, pooled normal |  | Reagent ordering |  |  | awaiting biotin labeling kits from Invitrogen
|  |  |  | pooled ANA etc. |  | Clinical samples availability |  |  | more samples can be obtained from SS and CL assays
|  |  |  | troubleshooting my Dab and buffers to obtain |  | Training with SS |  |  | experiments have been underway; troubleshooting my new Dab conjugates and buffers to compare response
comparable RLUs which SS had obtained
| Completed In |  |  |  |  | Completed In

## Heterophile Ab
| Experiment | Status | Timeline | Details
| Resource |  | Days
1 | MTP screen
2 | Conjugation
3 | Cross reactivity
4 | Interference
5 | Epitope map
6 | Theranos screen
7 | Training set
Antibody Pair finalized
8 | Plasma Screen
9 | Whole Blood Screen
10 | Capture titration
11 | Detection titration
12 | Detection Stabilizer
13 | Effect of Buffer
14 | Effect of Incubation time
15 | Reader protocol optimization
16 | Calibrator comparison
17 | Whole blood spike recovery
18 | Hematocrit effect
19 | Plasma spike recovery
20 | LLOQ/ULOQ
21 | Matrix effects
22 | Specificity: RF, HAMA
23 | Clinical samples
24 | Stability
Issues
1) Just got assigned with the assay (heterophilic antibody), doing literature review on the assay
2) Infectious mononucleosis is due to Epstein–Barr virus (EBV)
3) Commercially-available test kits are 70-92% sensitive and 96-100% specific
4) Tests are less sensitive in children or in adults during the early days of the illness (IgM Monospot test is + only for acute EBV Mononucleosis)
5) Test will generally not be positive during the 2-4 week incubation period before the onset of symptoms and will generally not be positive after active infection has subsided and hence the population with positive heterophile antibodies will be less
6) Especially in adolescents and young adults, the disease is characterized by fever, sore throat and fatigue, along with several other possible signs and symptoms; Can be mistaken for cytomegalovirus and toxoplasmosis;
7) According to the literature, it will be difficult to find Heterophile antibody positive even if they are EBV Positive.
8) There is no ELISA commerical kit available, commerial kits that are available are latex, RBC aggulutination and cartridge test kits.
9) Currently will be ordering reagents and waiting for some ordered reagents to arrive.
Completed In

## Completed_6-8-12
| CPT | Test Name | CMS Long Description | sample | Effect of SST | Sample processing | Other Notes | Clincal test pending | Siemens Validated in Li-Heparin | Siemens test notes
1 | 84443 | thyroid stim hormone | Thyroid stimulating hormone (TSH) | whole blood | no issues with serum gels |  |  |  | Yes
2 | 84153 | psa, total | Prostate specific antigen (PSA); total | whole blood | no issues with serum gels |  |  |  |  | PSA applicable to our assay rather than PSA3rd
3 | 84439 | thyroxine, free | Thyroxine; free | whole blood | no issues with serum gels |  |  |  | Yes
4 | 84702 | chorionic gonadotropin | Gonadotropin, chorionic (hCG); quantitative | whole blood |  |  |  |  | Yes
5 | 86140 | c-reactive protein | C-reactive protein; | whole blood |  |  |  |  | Yes
6 | 82728 | ferritin | Ferritin | whole blood | no issues with serum gels (agarose) |  |  |  | Yes
7 | 81025 | pregnancy , urine | Urine pregnancy test, by visual color comparison methods | urine |  |  | Need Clinical confirmation | Yes
8 | 86592 | syphilis non-trep qual | Syphilis test, non-treponemal antibody; qualitative (eg, VDRL, RPR, ART) | whole blood | no issues with serum gels | Serum to be collected witin 2 hrs of collection |  | Yes | Yes
9 | 86703 | hiv-1/hiv-2, single | Antibody; HIV-1 and HIV-2, single assay | whole blood |  |  |  | Yes
10 | 86618 | lyme disease antibody | Antibody; Borrelia burgdorferi (Lyme disease) | whole blood, csf | no issues with serum gels
11 | 84484 | troponin, quant | Troponin, quantitative | whole blood | no issues with serum gels
12 | 84436 | thyroxine, total | Thyroxine; total | whole blood | no issues with serum gels |  |  |  | Yes
13 | 86803 | hepatitis c ab | Hepatitis C antibody; | whole blood | no issues with serum gels |  |  | Yes
14 | 86038 | antinuclear antibodies | Antinuclear antibodies (ANA); | whole blood | no issues with serum gels
15 | 83001 | gonadotropin (fsh) | Gonadotropin; follicle stimulating hormone (FSH) | whole blood |  |  |  |  | Yes
16 | 84403 | total testosterone | Testosterone; total | whole blood | no issues with serum gels |  |  |  | Yes
17 | 86317* | rubella antibody, IgG | Antibody; rubella | whole blood | no issues with serum gels |  | Original assay list includes the CPT code of the IgM Rubella antibody -86762 |  | Yes
18 | 82670 | estradiol | Estradiol | whole blood | no issues with serum gels |  |  |  | Yes
19 | 84479 | thyroid (t3 or t4) | Thyroid hormone (T3 or T4) uptake or thyroid hormone binding ratio (THBR) | whole blood | no issues with serum gels
20 | 82553 | creatine, mb fraction | Creatine kinase (CK), (CPK); MB fraction only | whole blood | no issues with serum gels
21 | 84146 | prolactin | Prolactin | whole blood | no issues with serum gels |  |  |  | Yes
22 | 84480 | triiodothyronine (t3), Total | Triiodothyronine T3; total (TT-3) | whole blood | no issues with serum gels |  |  |  | Yes
23 | 83002 | gonadotropin (lh) | Gonadotropin; luteinizing hormone (LH) | whole blood
24 | 87430 | strep a ag, eia | Infectious agent antigen detection by enzyme immunoassay technique, qualitative or semiquantitative, multiple-step method; Streptococcus, group A | nasal | N/A | Throat swabs collected in culture medium |  | Yes
25 | 84144 | progesterone | Progesterone | whole blood |  |  |  |  | Yes
26 | 86787 | varicella-zoster antibody, IgG | Antibody; varicella-zoster | whole blood
27 | 84402 | testosterone, Free | Testosterone; free | whole blood | no issues with serum gels
28 | 82785 | ige | Gammaglobulin (immunoglobulin); IgE | whole blood
29 | 83036 | HbA1c
30 | 86235 | Nuclear Antigen Antibody | SSA
31 | 86235 | Nuclear Antigen Antibody | SSB
32 | 86704 | HBcAB, Total | Hepatitis B core antibody, Total |  |  |  |  | Yes
33 | 86376 | Microsomal Antibody | Thyroid Peroxidase antibody
34 | 86800 | Thyroglobulin antibody
35 | 86705 | HBcAB, IgM | Hepatitis B core antibody ,IgM |  |  |  |  | Yes
36 | 86665 | epstein-barr antibody
37 | 84481 | ft-3, free | Triiodothyronine T3; free
38 | 86430-1 | rheumatoid factor test - IgG | Rheumatoid factor; qualitative
39 | 86430-2 | rheumatoid factor test - IgA | Rheumatoid factor; qualitative
40 | 82306 | vitamin d, 25 hydroxy

## Test list_Development
| CPT | TestName | CMS Long Description | Under development | To start | Details
1 | 82607 | vitamin b-12 | Cyanocobalamin (Vitamin B-12); | √
2 | 87340 | hepatitis b surface ag, eia | Infectious agent antigen detection by enzyme immunoassay technique, qualitative or semiquantitative, multiple-step method; hepatitis B surface antigen (HBsAg) | √
3 | 82746 | folic acid serum, blood | Folic acid; serum | √
4 | 86235 -3 | nuclear antigen antibody - Sm |  | √
5 | 86235 -4 | nuclear antigen antibody - RNP |  |  | √ | resource limited
6 | 86235 -5 | nuclear antigen antibody - Scl-70 |  | √
7 | 86235 -6 | nuclear antigen antibody - Jo-1 |  | √
8 | 83970 | parathormone | Parathormone (parathyroid hormone) | √
9 | 86308 | heterophile antibodies | Heterophile antibodies; screening | √
10 | 82105 | alpha-fetoprotein, serum | Alpha-fetoprotein (AFP); serum | √
11 | 87400 -1 | influenza A, ag, eia | Infectious agent antigen detection by enzyme immunoassay technique, qualitative or semiquantitative, multiple-step method; Influenza, A or B, each | √
12 | 87400 -1 | influenza B, ag, eia |  | √
13 | 86431 | rheumatoid factor, quant | Rheumatoid factor; quantitative | √
14 | 84466 | transferrin | Transferrin | √
15 | 85379 | fibrin degradation, quant | Fibrin degradation products, D-dimer; quantitative | √
16 | 86709 | hep a antibody, igm | Hepatitis A antibody (HAAb); IgM antibody | √
18 | 80101-1 | ipecac | Reference method is LC/MS |  | √ | Reference assay is LC-MS based, have not seen ELISA based tests
19 | 80101-2 | ethyl glucuronide, | Drug screen, qualitative; single drug class method (eg, immunoassay, enzyme assay), each drug class | √
20 | 80101-3 | acid | Acid, LSD, LSD-25, lysergide, D-lysergic acid diethyl amide, N,N- diethyl- D- lysergamide |  | √ | Drug screen assay, need conjugates and/or binders
21 | 80101-4 | butorphanol | Stadol, Stadola |  | √ | Drug screen assay, need conjugates and/or binders
22 | 80101-5 | dolophine, | Methadone (also known as Symoron, Dolophine, Amidone, Methadose, Physeptone, Heptadon, Phy | √
23 | 80101-6 | flunitrazepam | Rohypnol, Rohipnola |  | √ | Drug screen assay, need conjugates and/or binders
24 | 80101-7 | heroin, | Diamorphine, Diacetylmorphine, Acetomorphine, (Dual) Acetylated morphine, Morphine diacetate |  | √ | Drug screen assay, need conjugates and/or binders
25 | 80101-8 | Cocaine |  |  | √ | Drug screen assay, need conjugates and/or binders
26 | 80101-9 | Codeine |  |  | √ | Drug screen assay, need conjugates and/or binders
27 | 80101-10 | Amphetamines |  |  | √ | Drug screen assay, need conjugates and/or binders
28 | 80101-11 | Phencyclidine (PCP) |  |  | √ | Drug screen assay, need conjugates and/or binders
29 | 80101-12 | Marijuana |  |  | √ | Drug screen assay, need conjugates and/or binders
30 | 80101-13 | Nalbuphine |  |  | √ | Drug screen assay, need conjugates and/or binders
Total

## Sheet1
Team
ELISA
1 | Sharada - Team Lead |  |  | R&D
2 |  | Xiaoyan Du |  | R&D
|  |  | * Mi Yan | R&D
3 |  | Andy Chen |  | R&D
4 |  | Tam Dang |  | R&D
5 |  | Amy Dinh |  | R&D
6 | Tina - Team Lead |  |  | R&D
7 |  | Evangelina |  | R&D
8 |  | Emma |  | R&D
9 |  | Sheena |  | R&D
10 | Ran Hu |  |  | R&D
11 | Nahal |  |  | R&D
12 | Ling Wang |  |  | Binders acceptance team
13 | Karen Shaw |  |  | Binders acceptance team
14 | Christopher Loi |  |  | Validation
MFG & Formulations
1 | Esther |  |  | Formulations
2 | Darren |  |  | Cartridge associate/ QC
3 | Samntha Zamora |  |  | Cartridge associate/ QC
4 | Salina Abusali |  |  | Cartridge associate/ Lab
5 | Javier |  |  | Coating/ Formulations
6 | Peter Low |  |  | Coating
7 | Andrew Kim |  |  | Coating/ Cartridge associate