## Assay list
| Assay |  |  |  | Assay Complete | In Development | Reagents identified | Next assignments | Other (date assigned, reagent status at assignment)
1 | thyroid stim hormone |  |  |  |  | Tina |  |  | 12/22/2010, Reagents ordered,
2 | glycosylated hemoglobin | [6 empty] | Andy ?
3 | vitamin d, 25 hydroxy |  |  |  |  | Tina |  |  | 1/20/2011, No reagents ordered.
4 | psa, total |  |  |  |  | Hong |  |  | 1/7/2011, most reagents ordered
5 | thyroxine, free
6 | chorionic gonadotropin
7 | c-reactive protein
8 | ferritin |  |  |  |  | Amy
9 | vitamin b-12 |  |  |  |  | Sharada |  |  | 1/24/2011, No reagents in-house
10 | hepatitis b surface ag, eia |  |  |  |  | Hong |  |  | 1/24/11, reagents ordered
11 | syphilis non-trep qual
12 | rbc antibody screen
13 | hiv-1/hiv-2, single Antibody; HIV-1 and HIV-2, single assay
14 | lyme disease antibody Antibody; Borrelia burgdorferi (Lyme disease | [6 empty] | Tina | 3/21/2011, No reagents ordered
15 | troponin T, quant |  |  |  |  | Sharada |  |  | 1/10/11, No reagents in-house
16 | thyroxine, total
17 | hepatitis c ab Hepatitis C antibody;
18 | folic acid serum, blood |  |  |  |  | Danielle |  |  | All reagents transferred on 2/10/11
19 | antinuclear antibodies | [6 empty] | Annapurna?
20 | gonadotropin (fsh)
21 | total testosterone
22 | rubella antibody
23 | estradiol
24 | thyroid (t3 or t4)
25 | creatine, mb fraction |  |  |  |  | Winny |  |  | 3/21/2011, No reagents in-house
26 | prolactin | [6 empty] | Xiaoyan
27 | triiodothyronine (t3) Triiodothyronine T3; total (TT-3)
28 | gonadotropin (lh)
29 | strep a w/optic | Infectious agent antigen detection by immunoassay with direct optical observation
30 | pregnancy , urine Urine pregnancy test, by visual color comparison methods
Other Assay Development
| Assay |  | Developer | Status | Issues
1 | IL-10 |  | Hong | Training set confirmation complete, Optimization | Original pair does not respond well to WHO standard

## TSH
Assay: | Thyroid Stimulating Hormone (TSH)
Developer: | Tina Noyes
Analyte Notes:
TSH is a glycoprotein hormone consisting of an alpha and beta protein subunit. The alpha subunit is identical to other pituitary hormones LH, FSH and hCG the beta subunit is unique and bestows specificity
Predicted Assay Challenges:
1. Due to the shared alpha subunit, cross reactivity with other pituitary hormones is a big concern, and pairs will need to be chosen on the basis of negligible cross reactivity
2. Normal patients will have mid range TSH while hyperthyroidism results in low TSH levels and hypothyroidism high TSH levels, therefore the disease range is both low and high.
3. Since TSH levels will be very low in hyperthyroidism whil other pituitary hormones may be high, only pairs with negligible cross raeactivity can be used
Date | Experiment | Notes
2/08 - 2/12/11 | Literature review and research on analyte and existing methods, ranges and relevant diseases
2011-02-08 | Test depleted serum, 2 TSH analytes, WHO standard, and clinical samples on the Alpco ELISA kit
2011-02-09 | Make serum calibrators for screening
2011-02-11 | quantify raw antibodies on nanodrop
2011-02-14 | Conjugate 22 antibodies as capture and detection
2011-02-15 | Quantify conjugates on nanodrop, Coat 22 microtitre plates for antibody screening of 462 pairs
2011-02-16 | Screen 462 pairs on 22 plates
2011-02-17 | Analyze screening data and rank pairs using manufacturer-supplied cross reactivity data and dose response
2011-02-22 | Coat MTPs for cross reacivity screening with FSH, LH and hCG
2011-02-23 | Run cross reactivity screening round 1 with best pairs from initial screen | All best pairs show unacceptabel criss reactivity
2011-02-23 | Coat MTPs for round 2 cross reactivity screening
2011-02-24 | Run cross reactivity screening round 2 | One pair yielded with no significant cross reacivity
2011-02-24 | Coat MTPs for round 3 cross reactivity screening
2011-02-25 | Run cross reacivity screening round 3 | 3 more pairs with no sigbnificant cross reactivity and suffucuent modulation to yield good sensitivity
2/26 to 2/27/11 | Screen best 4 pairs on Edison, simultaneous cross reactivity test (re-confirmed on Edison because surface and temperature differences make change Ab affinities) | Narrowed to 2 best pairs - some pairs with borderline cross reactivity in MTP were unaccepatble cross reacivity on Edison
2011-02-28 | Clinical traning set - run 11 clinical samples from across the range (already tested in kit) and check for correlation | Both best paris showed good correlation, slight favor to pair CAb12/DAb15
2011-03-02 | Matrix Screen - compare the matrix effects with the 2 best pairs | CAb12/DAb15 finalized as best pair based on clinical correlation, best sensitivity, and minimal matrix effects in whole blood, plasma, serum and buffer
3/3 to 3/8/11 | Tray QC for tips coated fails 3 successive runs - troubleshoot reagents - all MTP data looks fine, testing reagents individually on edison shows that tips must be producing the high CVs
Coating Tomtec is overdue for maintenance and showing aspiration problems, parts have been ordered and should be in place week of 03/21/11
2011-03-09 | Conjugate more of best pair antibodies, qualify on MTP
2011-03-10 | Progress on Edison on hold until coating Tomtec parts come in and are installed and tips can be successfully coated - switch to focus on laying groundwork for 25-OH Vitamin D assay, recruitiing etc
2011-03-16 | Epitope mapping by competition done with 18 different Ab's on MTP
Conclusion and Remaining experiments:
I am confident that the pair chosen will perform as required in all matrixes including whole blood and shows excellent correlation with predicate methods.
Now that the pair is finalized and bulk conjugates have been made (and 3mg of each Ab on order) as soon as tips can be produced again it will be a matter of days to finalize the optimization
Titrate Cab and Dab, test sample dilutions, test effect of incubation time, and post sample wash
Test clinical sample results on final conditions
While the coating tomtec is awaiting repair since all the steps left are optimization of the assay on the Edison.
So now I am focusing on research for 25-OH vit D and also now ordering reagents, clinical samples and doing literature review for the next assay - Lyme Disease Ab.

## 25-OH-Vit D
Assay: | 25-OH-Vitamin D (Calcidiol, 25-hydroxyvitamin D)
Developer: | Tina Noyes
Analyte Notes:
Calcidiol is a stable metabolite of cholecalciferol, used to determine Vitamin D status of the patient
Predicted Assay Challenges:
1. There are various metabolites and forms of vitamin D which lead to confusion in the nomenclature with manufacturers and suppliers
2. Calcidiol is bound to Vitamin D Binding Protein (VDBP) in the blood - most commerical methods require an organic extraction or protein precipition before analysis
3. Calcidiol is a small molecule and will require a competitive format and there are no commercially available conjugates of any kind
Research Notes: | I plan to test VDBP as a capture surface (binder) as a ready-made natural binder of this molecule. Specicificty will need to be tested
Capture antibodies will also be tested
Approach used with estradiol is likely to be beneficial - preincubation of sample with binder in solution to allow maximal binding, then introduce to capture surface
Have devised 3 different potential chemistries for conjugating 25-OH-vitamin D to ALP: requires purification steps and organic chemistry for which we do not have the equipment in place,
working with Pierce and Dojindo to obtain quotes and feasibility for having this contracted out to them
Date | Experiment | Notes
3/10/11 - present | Literature review and chemistry planning for conjugation and capture surfaces
2011-03-11 | Test clinical samples and analyte from EMD in Diazyme kit (colorimetric assay which requires protein precipitation step) | Kit performed poorly with its own controls and clinical samples. Method appears to be substandard
2011-03-15 | Test clinical samples and analyte from EMD in IDS kit (competitive EIA, no extraction) | Kit performed well and showed good correlation with reported values for the clinical samples - this kit can be used as a predicate method until analyzer is available
2011-03-16 | Qualify raw antibodies on nanodrop, conjugate anti-sheep, anti-mouse and anti-VDBP antibodies as capture surface
3/16 to 3/17 /11 | Test screen on MTP with 5 antibodies and leftover ELISa kit conjugate (calcidiol-biotin conjugate, use with strep-AP and Strep-HRP as second step) | Very weak dose response for 2 antibodies, not enough kit conjugate to test different levels and do checkerboard titration
Need to obtain concentrated analyte-ALP conjugate before testing of formats can proceed - Pierce conference call to be 3/23/11, Dojindo working on pilot test for the conjugation currently

## CTnT
Sharada
Round 2 of screening I did the following:
1. Ordered 20 more Abs majority of them from US biological, these have epitope and cross reactivity information.
2. For each pair a six point curve was tested, a dual screen: 3 point curve of the ARP native CTNT in assay buffer (purified from human heart tissue) at 2000, 500 and 0 pg/ml and a 3 level Biorad Liquichek cardiac serum controls. Per the datasheet the levels of cTNt in these are : 2320 pg/ml, 487 pg/mL and 107 pg/ml. I used the in house calibrators at comparable levels to get a sort of matched response. Hence the range used for this round of screening is higher , previous round the highest calibrator level was 100 pg/ml (based on the ARUP lab reference range).
3. The serum cardiac controls however also include Troponin C.
4. It might be in our favor to increase the range to see if we can get “any” response.
5. For the assay buffer calibrators I made in bulk , aliquoted and flash froze and used the same lot for the entire screening.
6. Similar approach was taken with the serum controls.
7. Note: The datasheet for the controls indicated poor stability for the CTNT in serum and indicated that these would have to be thawed and aliquoted immediately for future use. Literature references (emailed to you) do not indicate any issues with stability if cTNT in whole blood or serum , so this must be specific to these samples we do not know how old these are. Nevertheless it was a good idea to thaw them once and aliquot them en mass.
8. Note again: There was only 12 ml of each level avaible (it is expensive) hence in the screen the analyte volume was decreased to 20 uL for all calibrators. Additionally the serum controls were diluted 1:2 in BB.
Although I still have to finish all of the screening (5 of the 17 Abs from USBiol were backordered and were supposed to arrive on 3/15 but only arrived today, plus I was shorted out on about 6 Abs and they have not replaced them yet despite several emails.) the results after looking at 760 pairs is encouraging. Winners were selected based on S/B of >15 for the native analyte AB curve since the Troponin C in the serum might give us false positives if the pairs have cross reactivity. In all these cases there was also similar or higher H/L values for the serum controls.
9. There are ~50 pairs that have a S/B >15.
10. Further weeding out of winners was done based on Std2/Std 3 ratios (in this case CTNT500/cTNT0) to see if any of these pairs can be picked based on the sensitivity they “might” afford. A final count of 6 pairs was picked. All these have Std2/Std3 ratios >10.
11. It remains to be seen how much CR I will encounter.
Next:
1. I am running low on the conjugates for these pairs and also the raw abs so I am planning to use only 5 ug/mL of Capture ab for the CR and Interference experiments,
2. I am ordering 1 mg of these Abs (some of them are common) because it takes a very long time (~ month) to receive stuff from US BIOLOgical.
3. Analytes for CR test (all assay buffer stds) : control(native cTNT ARP), cTNI, cTNC, skeletal TNT, Troponin complex (cardiac T, I and C isoforms) in a 1:1:1 mix, along with the Biorad cardiac controls.
Detection Ab
|  | #1 | #2 | #3 | #4 | #5 | #6 | #7 | # 8 | # 9 | #10 | #11 | #12 | #13 | #14 | #15 | #16 | #17 | #18 | #20 | #21 | #22 | #25 | #34 | #35 | #36 | #38 | #39 | #40 | #23 | #24 | #26 | #27 | #28 | #29 | #30 | #31 | #32 | #33 | #37
Capture Ab | #1
#2
#3
#4
#5
#6
#7
| #8 | Detection Ab #5: ROUND 1 SCREENING
#9
#10
#11
#12
#13
#14
#15
#16
#17
#18
#20
#21
#22
#23
#25
#28
#29
#32
#33
#34
#35
#36
#38
#39
#40
#24
#26
#27
#30
#31
#37
Ab pair screening
Date | 2011-03-11
Expt | Ab pair screening for cardiac troponin T assay: Each pair was tested with a 6 point calibrator curve: 3 levels of the native cardiac Troponin T analyte High (2000), Medium (500) and 0 pg/ml
and 3 levels of the Biorad Liquichek Cardiac Controls , High(2320 pg/mL), Medium (484 pg/ml) and low (107 pg/ml)
MTP | Nunc, Assay plate 384 well, white, Maxisorp, Lot# 1008780, Cat# 460372
Protocol | 30'UA/30'Cab/3x wash/ 10' fixative
UA | 20 ug/mL in carbonate bi carbonate buffer, stock 10 mg/ml, Leinco, P# A110, Lot# 1210L200
Cab | 10 ug/mL in blocking buffer, See attached sheet for details on Ab conjugations
Dab | Dab in StabilZyme 100 ng/mL, See attached sheet for details on Ab conjugations
Buffers | Blocking buffer (3% BSA, 0.05% NaN3 in TBS), Fixative buffer (3% BSA, 0.05% NaN3, 0.4% Sucrose, 0.01% Glycerol, in TBS) DOM 1/28/11 filtered
Analyte | Cardiac troponin T, purified from human heart tissue, American research products ARP, Cat# 05-52095, Lot# 10061
Serum controls | Biorad, Liquichek, Cardiac Markers plus control, Cat# 180X, Trilevel MiniPak , 3x3 mL, | Detection Ab #10: Level 1 (29761) is 0.107 ng/mL
Assay | 10/10/10, 3x wash after Dab | Detection Ab #10: Level 2 (29762) is 0.484 ng/mL
Substrate | PhosphaGLO AP substrate, Cat# 55-60-04, Lot #100696 | Detection Ab #10: Level 3 (29763) is 2.32 ng/mL
Dilution | CTNT AB calibrators none
| Serum controls | Detection Ab #2: Diluted 1:2 in BB

## IL-10
Assay: | IL-10
Assay Developer: | Hong Li
Experiment complete date |  | Experiment | Information desired | Comments
10/14/2010-10/22/2010 | 1 | MTP screen | Summary information of antibody screens, 4 point curves | Show data on all pairs, as well as summary table with color coded response
| a. | 19 antibodies conjugation
| b. | Coating 21 plates
| c. | Assay 21 plates
10/25/2010-10/27/2010 | 3 | Cross reactivity | For 4 best Ab. Pairs:Test cross reactivity with related analytes, 6 point standard curves | On MTP, Range of the cross reactants tested will come from commercial kits or Literature based on Clinical disease and normal healthy levels
| a. | Coating 4 plates
| b. | Assay 4 plates
10/25/2010-10/27/2010 | 4 | Interference | For 4 best Ab. Pairs:Test interference with related analytes, 6 point standard curves | On MTP, Effect of related analyte at 3x the highest level on assay
| a. | Coating 4 plates
| b. | Assay 4 plates | For 4 best Ab. Pairs | On MTP, Optional
10/28/2010-10/29/2010 | 6 | Theranos screen | For 3 best Ab. Pairs: Summary information for 6 point curves in A.B, whole blood and plasma | On readers
11/1/2010-11/2/2010 | 7 | Training set | For 1 Ab. Pair correlation of N=10 clinical samples spanning the assay range range | Confirmation that antibody pair will recognize Clinical samples
|  | Antibody Pair finalized | Based on the best dose response and % recovery of above experiments
2010-11-03 | 8 | Plasma Screen | Screen normal plasma to see ranges | Minimum 10 samples screened
11/23/28/2010 | 9 | Whole Blood Screen | Screen normal blood to see ranges. | Minimum 10 samples screened
11//2/2010 | 10 | Capture titration | Show table with selected level and higher/lower levels | Titrate at 20, 10, 5, 2.5 ug/mL
11//4/2010 | 11 | Detection titration | Show table with selected level and higher/lower levels | Titrate at 4 levels
2011-11-23 | 13 | Effect of Buffer | Test various blocking buffers for coating and analyte diluent compositions
11/22, 11/24 | 15 | Edison protocol optimization | Different diliution effects, impact of PSW
2010-12-07 | 16 | % CV | Mid range analyte on 24 instruments | HT tomtec coating of 24 trays, Mean % CV and tray differential should not be >15%
2010-12-08 | 17 | Precision tests | 3 lot inter and intra assay results with standard curve and 5 QC samples, 3 cartridges per lot | Standard curve + 5 QC levels at LLOQ, Low QC, Mid QC and High QC,) Need %CV on back calculated concentrations, use as many machines as available
2010-11-30 | 18 | Calibrator comparison | Run our controls on reference kit and vice versa, calibrate with WHO standards | To calibrate our standards
2010-12-09 | 19 | Dilution linearity | Serial dilution of high samples in buffer and sera | low endogenous plasma/ sera needed, in its absence high clinical sample can be diluted with low sample
2010-12-11 | 20 | Whole blood spike recovery | Spike into blood with known (low) endogenous level at 8 point curve | show calibration of tips as well as calculated concentrations
2010-12-11 | 21 | Hematocrit effect | Plasma from spiked blood tested, plot W.B vs. plasma
2010-12-11 | 22 | Plasma spike recovery | Spike into plasma with known (low) endogenous level at 8 point curve | show calibration of tips as well as calculated concentrations
2010-12-12 | 23 | LLOQ/ULOQ | Use 2x above ULOQ and below LLOQ to define assay limits | minimum 8-point standard curve in matrix
2011-01-27 | 24 | Selectivity | Spike into W.B at 3 levels over endogenous to calculate recovery | N=10 whole blood samples, 5 male and 5 female
2010-12-16 | 25 | Stability | CAB, DAB stability, Reference tips in some cases, with a 4-point standard curve | At 4ºC and RT for 24 weeks
2010-12-09 | 26 | Matrix effects | Spike recovery in hemolysed or lipemic samples for 6 spike levels | % recovery back calculated from a standard curve
2010-12-10 | 27 | Extended range | Test linearity and dose reposnse of standard curve at 4x of ULOQ
| 28 | Clinical samples | N=20, spanning the range to show correlation with kit (See Traing set section) | Calibration with 12 point curve with anchor points at 2x higher than ULOQ and 2x lower than LLOQ
Additional work from list above:
2010-11-21 | 1. Effect of other Dab (IL-10 Dab + IL-12 Dab + IFN-γ Dab) on IL-10 assay w/5X dilution protocol _
2010-11-22 | 2. 10X dilution protocol: ABA2_ver2 svn-5311_ assay w/ IL-10 Dab + IL-12 Dab + IFN-γ Dab
2010-11-23 | 3. 10X dilution protocol with additional wash: ABA2_ver3 svn-5318_ assay w/ IL-10 Dab + IL-12 Dab + IFN-γ Dab
2010-11-30 | 4. Test ABA_2 serum mixed calibrators on Edison (IL-10, IL-12, IFN-γ)
2010-01-17 | 5. ABA_2 calibration
6. IL-10 assay troubleshooting:
2010-11-14 | 1) IL-10 analyte comparison (Run on R&D ELISA Kit)
2010-11-15 | 2) IL-10 analyte comparison (Run on Edison)
2010-12-21 | 3) IL-10 analyte comparison (Run on MTP, Theranos assay)
2010-12-21 | 4) IL-10 analyte comparison with alternative antibody pairs (Run on Edison)
Cab 16 / Dab 11
Cab 8 / Dab 10
2010-12-28 | 5) IL-10 analyte comparison with Cab 11/ Mixed Dab (Dab 8, Dab9, Dab10, Dab11 all together) run on Edison
6) IL-10 analyte comparison run on Abcam IL-10 ELISA kit
2011-01-19 | 7) Effect of N-terminal Met of IL-10
R&D IL-10 analyte (Cat# 217-IL) we used for IL-10 assay is: Baculavirus derived , without N-terminal Met
R&D IL-10 analyte (Cat# 1064-IL) is: E. coli-devived, with N-terminal Met
2/3/, 2/4, 2/8, 2/9, /2/10 | 8) Antibody screening with WHO IL-10 analyte
Total screened antibody pairs: 18
Total tested antibody combination: 324
2011-02-14 | 9) IL-10 (WHO IL-10 Standard) spike recovery with Cab4/Dab3 (Best antibody pair from antibody screening with WHO IL-10 analyte)
Spike into whole blood
Plasma from spiked whole blood
Spike into plasma
| IL-10 blocking buffer standard curve was used for Il-10 calibration |  |  |  | ``
Out of antibody
2011-02-16 | 10) WHO IL-10 serum standard curve & plasma samples screening (total 6 plasma samples from normal subjects)
2011-03-02 | 11) Clinical samples screening on ELISA kit (total 29 samples)
2011-03-02 | 12) Clinical samples correlation with Cab4/Dab3
2011-03-03 | 13) Clinical samples correlation with Cab11/Dab9
2011-03-08 | 14) Spiked WHO IL-10 samples (3 different level of IL-10 with 5 different plasma samples ) correlation with Cab4/Dab3
Spiked WHO IL-10 samples (3 different level of IL-10 with 5 different plasma samples ) with Cab4/Dab3_on Edison
Spiked WHO IL-10 samples (3 different level of IL-10 with 5 different plasma samples ) with Cab4/Dab3_on R&D ELISA kit
2011-03-10 | 15) IL-10 analyte comparison with Cab4/Dab3 on Edison
WHO IL-10 (E. coli-devived)
R&D analyte (Cat# 217-IL, used for IL-10 assay Cab11/ Dab9, Baculavirus derived )
R&D analyte (Cat# 1064-IL , I ordered for assay thouble shooting, E. coli-devived)
2011-03-11 | 16) IL-10 analyte comparison with Cab4/Dab3 on R&D ELISA kit & on Edison (ELISA kit assay diluent was used for all those three analyte)
ELISA kit IL-10 analyte
WHO IL-10 (E. coli-devived)
R&D analyte (Cat# 1064-IL , I ordered for assay trouble shooting, E. coli-devived)
2011-03-11 | 17) Protocol quick test: Co incubation protocol vs. 10/10/10 protocol
2011-03-16 | 18) Cross reactivity test with antibody pair Cab4/Dab3
Tested analyte: Il-4, IL-6, IL-8, IL-12, IFNg, TNFa
2011-03-17 | 19) Interference test with antibody pair Cab4/Dab3
Tested analyte: Il-4, IL-6, IL-8, IL-12, IFNg, TNFa

## PSA
| Assay: |  | Total PSA
| Assay Developer: |  | Hong Li
Date | Complete |  | Experiment | Information desired | Comments
2/3/2011 (7 days) |  | 1 | Antibody screen1 | Summary information of antibody screens, 4 point curves | Show data on all pairs, as well as summary table with color coded response
2/15/2011 (2 days) |  | 2 | Antibody screen2
| NA | 3 | Conjugation Type: SH or NH2 | Compare SH vs. NH2 conjugations (For assays with limited antibodies to screen) | MTP
2011-02-23 |  | 4 | Cross reactivity1 | For 4 best Ab. Pairs:Test cross reactivity with related analytes, 6 point standard curves | On MTP, Range of the cross reactants tested will come from commercial kits or Literature based on Clinical disease and normal healthy levels
2011-03-18 |  | 5 | Cross reactivity2
|  | 6 | Interference | For 4 best Ab. Pairs:Test interference with related analytes, 6 point standard curves | On MTP, Effect of related analyte at 3x the highest level on assay
| NA | 7 | Epitope map | For 4 best Ab. Pairs | On MTP,
2011-02-28 |  | 8 | Theranos screen | For 3 best Ab. Pairs: Summary information for 6 point curves in A.B, whole blood and plasma | On readers
Order 500 ug each of best 3 antibody pairs
|  | 9 | Training set | For 1 Ab. Pair correlation of N=15 clinical samples spanning the assay range range | Confirmation that antibody pair will recognize Clinical samples
2011-02-23 |  | 10 | Calibrator comparison | Run our controls on reference kit and vice versa, calibrate with WHO standards | To calibrate our standards
Antibody Pair finalized; Order 3 mg of final pair
|  | 11 | Plasma Screen | Screen normal plasma to see ranges | Minimum 10 samples screened
|  | 12 | Whole Blood Screen | Screen normal blood to see ranges. | Minimum 10 samples screened
|  | 13 | Capture titration | Show table with selected level and higher/lower levels | Titrate at 20, 10, 5, 2.5 ug/mL
|  | 14 | Detection titration | Show table with selected level and higher/lower levels | Titrate at 4 levels
| NA | 15 | Detection Stabilizer | Biostab vs. Stabilzyme | Can be combined with Titration
|  | 16 | Effect of Buffer | Test various blocking buffers for coating and analyte diluent compositions
Order MFG lot of Tips for the remaining experiments
|  | 17 | Sample dilution | Test different dilutions and impact on assay sensitivity | e.g if 1:5 is optimal for assay test 1:5, 1:10, 1:15 (Hand dilutions OK if Protocols are unavailable)
|  | 18 | Whole blood spike recovery | Spike into blood with known (low) endogenous level at 8 point curve | show calibration of tips as well as calculated concentrations
|  | 19 | Hematocrit effect | Plasma from spiked blood tested, plot W.B vs. plasma
|  | 20 | Plasma spike recovery | Spike into plasma with known (low) endogenous level at 8 point curve | show calibration of tips as well as calculated concentrations
|  | 21 | Stability | CAB, DAB stability, Reference tips, with a 4-point standard curve | At 4ºC and RT for 24 weeks, CAB as Tips, DAB on MTP
|  | 22 | Matrix effects | Spike recovery in hemolysed, lipemic samples for 6 spike levels | % recovery back calculated from a standard curve
|  | 23 | Selectivity (Optional) | If Whole blood spike recovery does not show acceptable % recovery | N=10 whole blood samples, 5 male and 5 female

## HCG_urine
| Assay: |  | hCG
| Assay Developer: |  | Hong Li
| Complete |  | Experiment | Information desired
2011-01-17 |  | 1 | Urine Matrix effects | Good dose response was seen for the hCG assay in Urine. Modulation was comparable to Plasma.

## Ferritin
Amy Dinh
Assay: | Ferritin
Date | Experiment | Notes
1/17/11 to 1/26/11 | Literature review and research on analyte and existing methods, ranges and relevant diseases
2011-02-15 | Test 20 clinical samples on the Alpco ELISA Ferritin kit
Based on kit's data, I have a good range of Ferritin levels across the samples so OK to proceed forward with assay development, keeping in mind that con-currently I also balance with QC work and LH assay.

## LH
Amy
Assay: | LH
Developer: | Transferred from ET to AD in February 2011
Date | Experiment | Notes
2/17-2/25 | Clinical traning set - run total of 30 clinical samples from across the range and check for correlation
3//22/11 | Stability: DAB stability with a 4-point standard curve. At 4ºC and RT for 24 weeks, DAB on MTP
This week | Plasma Screen: Screen normal plasma to see ranges, minimum 10 samples screened
Whole Blood Screen: Screen normal blood to see ranges, mminimum 10 samples screened
| Whole blood spike recovery. Spike into blood with known (low) endogenous level at 8 point curve. Show calibration of tips as well as calculated concentrations |  | show calibration of tips as well as calculated concentrations
Hematocrit effect. Plasma from spiked blood tested, plot W.B vs. plasma.
| Plasma spike recovery. Spike into plasma with known (low) endogenous level at 8 point curve. Show calibration of tips as well as calculated concentrations |  | show calibration of tips as well as calculated concentrations

## QC
Amy
March 11 to March 18, 2011
Date | Test
2011-03-11 | Substrate Luminescene Testing (pre-built cartridges from Mfg.): DOF 3/10/11
2011-03-11 | ACE-011 Cartridge Stability Study: Time Point = Week 24
2011-03-14 | Testing Newly Formulated TNFα, IL-6, and IL-8 Cab Cojugates for ABA1 Cartridge Stability Study, N=3 cartridges for a selected 9-point curve)
2011-03-15 | Testing Newly Formulated TNFα and IL-8 Dab Cojugates for ABA1 Cartridge Stability Study , N=3 cartridges for a selected 9-point curve)