## Male Health Layout
Part Number:
Cartridge Name: Male Health
Cartridge ID: 118
Assay Protocol: Male_Health svn 6907
NOTE: Requires 35uL of whole blood added to cartridge
(Row C) Tip Placement:
Column | Assay | Test/Ref | In Tip Ref Conc | In Sample Ref Conc | Units
1 | Ferritin | Test | -- | -- | ng/mL
2 | HbA1c | Test | -- | -- | mg/mL
3 | TSH | Test | -- | -- | uIU/mL
4 | tPSA | Test | -- | -- | ng/mL
5 | 25-HydroxyvitaminD | Test | -- | -- | ng/mL
6 | no not use | -- | -- | -- | --
7 | no not use | -- | -- | -- | --
8 | no not use | -- | -- | -- | --
(Row D) Detection Ab in Cartridge:
Position | Reagent | Conc |  | Volume (uL)
F3 | Blocking Buffer | 3% BSA in TBS |  | 400
F4 | 25OHD CAb | Ab4 100 ng/mL and Ab2 50 ng/mL in Super Block |  | 60
F5 | 25OHD-AP | Dojindo 25-HydroxyvitaminD3-AP 1:10,000 in Stabilzyme AP |  | 60
F7 | 25OHD Lyophilized Pepsin Tube | Pepsin 480ug (lyophilized)
D7 | 25OHD Neutralization Buffer | 50mM TRIS buffer pH 8.0 (not TBS) |  | 200
D8 | 25OHD Extraction Buffer | 100mM Citrate Buffer pH 2.6 |  | 100
B8 | HbA1c Extraction Buffer | See formulation |  | 300
A6 | tPSA/TSH DAb Mixture 2x | 50 ng/mL tPSA DAb and 20 ng/mL TSH DAb in Stabilzyme AP |  | 60
A7 | Ferritin DAb 2x | 200 ng/mL in Stabilzyme AP |  | 60
A8 | HbA1c DAb 10x | 250 ng/mL in House Stabilizer |  | 60
C1-6 | Wash Buffer | - |  | 400
A1-5 | Substrate | - |  | 100
E7 and E8 | Blood Tips (2) | -
A1 - A5 | KPL PhosphaGlo Substrate | -- |  | 80
B1 - B6 | Enzo 1x Wash Buffer | -- |  | 400
Component Pick List:
Part No. | Description |  | Quantity
Note: Quantity per component is based on single cartridge requirements.
Refer to Work Order for total required cartridge quantity.
1
2
3
4
5
6
7
8
O
1 2 3 4 5 6 7 8
A
B
C
D
E
F

## Ferritin Calibration
Ferritin Calibration
6-26-13 SZ
Female Health Demo Cartridges
Protocol: Female_Health svn
Calibrators: Ferritin Calibrators prepared by Formulation Group DOM: 9-25-12
Conclusion
1. The modulation is tracking with historical data and the CVs are below 20% except for tray 2.
2. released trays:
Ferritin Calibration | [12 empty] | Data |  | Model Predictions |  | Model Type | LogLin 4PL | [6 empty] | Date | 2013-06-26
[9 empty] | Back Calculation |  |  |  | Conc | RLU | RLU → Conc | Conc → RLU | Model Equation | log10(RLU) = b1+(b2-b1)/(1+(Conc/b3)^b4 | [6 empty] | Name | SZ
Calibrator # | Concentration ng/ml | [Ferritin] re-assigned per CLIA, ng/ml | Barcode | Edison # | Tip 1 | Mean | CV | Modulation | Tip 1 | Mean | CV% | %Recovery | 0 | 235.2534 | OORL | 192.6152 | Calibration Equation | If RLU <= SignalMax & RLU >= SignalMin Then Conc = b3 * (((b2 - b1) / (log10(RLU) - b1)) - 1) ^ (1 / b4) End If | [6 empty] | Cartridge Lot | 1234
1 | 1500 | 1500 | 916245643802300121 | e000169 | 249057.8721 | 257006.9388 | 4.374% | 1291.0259 | 1503.2831 | 1556.0619 | 4.797% | 103.7% | 0 | 224.0724 | OORL | 192.6152 | [8 empty] | Analyte | Ferritin
|  |  | 916245643802300122 cartridge did not finish protocol | [9 empty] | 0 | 137.8899 | OORL | 192.6152 | [8 empty] | Units | ng/ml
|  |  | 916245643802300123 | e000268 | 264956.0054 |  |  |  | 1608.8407 |  |  |  | 5.84 | 1487.3215 | OORL | 1375.1968 | [8 empty] | Cartridge Id | 118
2 | 1000 | 1015 | 916245643802300124 | e000201 | 150264.6566 | 125196.2857 | 28.32% | 628.8999 | 883.0394 | 600.3761 | 46.12% | 59.15% | 5.84 | 1382.6823 | OORL | 1375.1968 | [8 empty] | Dilution | 10000
|  |  | 916245643802300125 | e000187 | 100127.9148 |  |  |  | 588.4365 |  |  |  | 5.84 | 1269.4981 | OORL | 1375.1968 | [8 empty] | Reagent Lot | 1234
|  |  | 916245643802300126 | e000199 | 54931 |  |  |  | 329.6523 |  |  |  | 12.2 | 2043.3125 | 10.3741 | 2307.2013
3 | 500 | 519 | 916245643802300127 | e000231 | 109237.2125 | 98744.6341 | 15.03% | 496.025 | 641.1757 | 580.6772 | 14.72% | 111.9% | 12.2 | 2008.955 | 10.1372 | 2307.2013 | [8 empty] | desired LLOQ | 12.2
|  |  | 916245643802300128 | e000355 | 81764.8761 |  |  |  | 482.9065 |  |  |  | 12.2 | 2200.5475 | 11.4609 | 2307.2013 | [8 empty] | desired ULOQ | 1500
|  |  | 916245643802300129 | e000230 | 105231.8136 |  |  |  | 617.9494 |  |  |  | 29.5 | 5088.309 | 31.4168 | 4807.5004 | [8 empty] | max LLOQ | 12.2
4 | 250 | 280 | 916245643802300130 | e000280 | 49923.3368 | 41902.6012 | 17.76% | 210.4898 | 301.0018 | 254.8212 | 16.82% | 91.01% | 29.5 | 5220.7847 | 32.3188 | 4807.5004 | LLOQ | 12.2 | ng/ml |  | Model Parameters |  | SE |  | min ULOQ | 1500
|  |  | 916245643802300131 | e000049 | 40557.6529 |  |  |  | 247.1697 |  |  |  | 60.2 | 11047.1981 | 70.6639 | 9422.0753 | ULOQ | 1500 | ng/ml |  | b1 | 2.2847 | 0.044459
|  |  | 916245643802300132 | e000316 | 35226.814 |  |  |  | 216.2922 |  |  |  | 60.2 | 12027.2849 | 76.9046 | 9422.0753 | desired LLOQ | 12.2 | ng/ml |  | b2 | 6.8482 | 0.45245 |  | Data Grouping | by tip
5 | 100 | 104 | 916245643802300133 | e000339 | 29700 | 13947.3138 | 19.66% | 70.0617 | 183.9969 | 120.6117 | 46.61% | 116% | 60.2 | 11588.4899 | 74.1167 | 9422.0753 | desired ULOQ | 1500 | ng/ml |  | b3 | 225.7309 | 120.3159
|  |  | 916245643802300134 | e000115 | 12008.0315 |  |  |  | 76.7825 |  |  |  | 104 | 12008.0315 | 76.7825 | 16363.5422 | LLOD | Not Reported | ng/ml |  | b4 | -0.40201 | 0.03749 |  | Precision criteria for estimation of LLOQ | On
|  |  | 916245643802300135 | e000071 | 15886.5961 |  |  |  | 101.0556 |  |  |  | 104 | 15886.5961 | 101.0556 | 16363.5422 | LLOQ accuracy | 87.3558 | % |  | b5 | #N/A | #N/A
6 | 50 | 60.2 | 916245643802300136 | e000056 | 11047.1981 | 11554.3243 | 4.249% | 58.041 | 70.6639 | 73.8951 | 4.231% | 122.7% | 280 | 40557.6529 | 247.1697 | 46261.1267 | LLOQ precision | 6.6228 | %
|  |  | 916245643802300138 | e000265 | 12027.2849 |  |  |  | 76.9046 |  |  |  | 280 | 35226.814 | 216.2922 | 46261.1267 | Average Residuals | 13.8928 | % |  | 0 |  |  |  | Preferred models (if any) | 4PL
|  |  | 916245643802300139 | e000140 | 11588.4899 |  |  |  | 74.1167 |  |  |  | 519 | 109237.2125 | 641.1757 | 88062.9069 | Error in prediction: Best case | 14.2083 | %
6 | 25 | 29.5 | 916245643802300140 | e000111 | 5088.309 | 5154.5469 | 1.817% | 25.8929 | 31.4168 | 37.5423 | 26.21% | 127.3% | 519 | 81764.8761 | 482.9065 | 88062.9069 | Error in prediction: Expected | 16.1456 | %
|  |  | 916245643802300142 | e000165 | 7694 |  |  |  | 48.8914 |  |  |  | 519 | 105231.8136 | 617.9494 | 88062.9069
|  |  | 916245643802300143 | e000102 | 5220.7847 |  |  |  | 32.3188 |  |  |  | 1015 | 150264.6566 | 883.0394 | 172071.0437 | Signal Min | 1954.1633 | RLU | [5 empty] | fingerstickmultiplier | 1.6
6 | 10 | 12.2 | 916245643802300144 | e000122 | 2043.3125 | 2084.2717 | 4.901% | 10.4699 | 10.3741 | 10.6574 | 6.623% | 87.36% | 1500 | 249057.8721 | 1503.2831 | 248559.2924 | Signal Max | 282159.6666 | RLU | [5 empty] | venousmultiplier | 1.6
|  |  | 916245643802300145 | e000308 | 2008.955 |  |  |  | 10.1372 |  |  |  | 1500 | 264956.0054 | 1608.8407 | 248559.2924 | [6 empty] | 4PL |  | plasmamultiplier | 1.6
|  |  | 916245643802300146 | e000028 | 2200.5475 |  |  |  | 11.4609 | [7 empty] | Outliers detected based on within-standard variance and manual detection | Outliers |  |  |  | ULOQ accuracy | 103.7375 | % | srerun_lo | 50
6 | 5 | 5.84 | 916245643802300147 | e000268 | 1487.3215 | 1379.834 | 7.895% | 6.9313 | 6.5858 | 5.8738 | 12.26% | 100.6% | [5 empty] | Conc | RLU | 1 |  | ULOQ precision | 4.7968 | % | srerun_hi | 1000000
|  |  | 916245643802300148 | e000339 | 1382.6823 |  |  |  | 5.8895 | [7 empty] | Statistically Detected ■ Manually Selected ■ | 29.5 | 7694 | 1
|  |  | 916245643802300149 | e000280 | 1269.4981 |  |  |  | 5.146 | [8 empty] | 104 | 29700 | 0
6 | 0 | 0 | 916245643802300150 | e000140 | 235.2534 | 199.0719 | 26.76% | 1 | 0.012431 | [8 empty] | 280 | 49923.3368 | 0
|  |  | 916245643802300151 | e000187 | 224.0724 |  |  |  | 0.0061367 | [8 empty] | 1015 | 54931 | 0
|  |  | 916245643802300152 | e000115 | 137.8899 |  |  |  | #NUM! | [8 empty] | 1015 | 100127.9148 | 0
[18 empty] | 1500 | 1 | 0

## HbA1c Calibration
HbA1c Calibration
6-28-13 SZ
Male Health Demo Cartridges
Protocol: Male_Health svn
Calibrators: HbA1c Whole Blood Calibrators DOM: 2-10-12 TN
Conclusion
1.
2. released status:
HbA1c Calibration | [11 empty] | Data |  | Model Predictions |  | Model Type | LogLin Quadratic | [6 empty] | Date | 2013-07-01
[8 empty] | Back Calculation |  |  |  | Conc | RLU | RLU → Conc | Conc → RLU | Model Equation | log10(RLU) = b1 + b2 Conc + b3 Conc^2 | [6 empty] | Name | SZ
Calibrator # | Assigned [HbA1c] mg/mL | Barcode | Edison # | Tip 2 | Mean | CV | Modulation | Tip 2 | Mean | CV% | %Recovery | 0 | 151.9366 | OORL | 1996.1086 | Calibration Equation | If RLU <= SignalMax & RLU >= SignalMin Then If b3 > 0 Then Conc = Max( (-b2 + sqrt(b2 ^ 2 - 4 * b3 * (b1 - log10(RLU)))) / (2 * b3), (-b2 - sqrt(b2 ^ 2 - 4 * b3 * (b1 - log10(RLU)))) / (2 * b3)) End If If b3 > 0 Then Conc = Min( (-b2 + sqrt(b2 ^ 2 - 4 * b3 * (b1 - log10(RLU)))) / (2 * b3), (-b2 - sqrt(b2 ^ 2 - 4 * b3 * (b1 - log10(RLU)))) / (2 * b3)) End If End If | [6 empty] | Cartridge Lot | 1234
1 | 17.7 | 916245643802300217 | e000115 | 28157 | 53118.6257 | 24.04% | 395.4963 |  |  |  |  | 0 | 152.6582 | OORL | 1996.1086 | [8 empty] | Analyte | HbA1c
|  | 916245643802300218 | e000339 | 62148 | [7 empty] | 3.2 | 4506.8028 | 2.7698 | 5065.2187 | [8 empty] | Units | mg/ml
|  | 916245643802300219 | e000280 | 44089.2514 | [7 empty] | 3.2 | 6250.5073 | 3.9999 | 5065.2187 | [8 empty] | Cartridge Id | 118
2 | 11.1 | 916245643802300220 | e000268 | 35216.4325 | 34441.5599 | 3.182% | 256.4357 |  |  | #DIV/0! |  | 6.3 | 12078.4395 | 6.7707 | 10882.1776 | [8 empty] | Dilution | 10000
|  | 916245643802300221 pump error | [9 empty] | 6.3 | 10088.8418 | 5.9669 | 10882.1776 | [8 empty] | Reagent Lot | 1234
|  | 916245643802300222 | e000187 | 33666.6873 | [7 empty] | 6.3 | 13195.671 | 7.1815 | 10882.1776
3 | 9.2 | 916245643802300223 | e000201 | 27 | 21937.9619 | 11.5% | 163.3397 |  |  | #DIV/0! | 0% | 8 | 11288.2465 | 6.4637 | 15628.5393 | [8 empty] | desired LLOQ | 3.2
|  | 916245643802300224 | e000071 | 20153.7896 | [7 empty] | 8 | 14083.8698 | 7.4912 | 15628.5393 | [8 empty] | desired ULOQ | 17.7
|  | 916245643802300225 | e000140 | 23722.1342 | [7 empty] | 9.2 | 20153.7896 | 9.327 | 19690.3995 | [8 empty] | max LLOQ | 3.2
4 | 8 | 916245643802300226 | e000355 | 11288.2465 | 14973.5231 | 28.06% | 111.4858 |  |  | #DIV/0! | 0% | 9.2 | 23722.1342 | 10.2544 | 19690.3995 | LLOQ | 3.2 | mg/ml |  | Model Parameters |  | SE |  | min ULOQ | 17.7
|  | 916245643802300227 | e000095 | 14083.8698 | [7 empty] | 11.1 | 35216.4325 | 12.8739 | 27235.0289 | ULOQ | 17.7 | mg/ml |  | b1 | 3.3002 | 0.13794
|  | 916245643802300228 | e000230 | 19548.4531 | [7 empty] | 11.1 | 33666.6873 | 12.54 | 27235.0289 | desired LLOQ | 3.2 | mg/ml |  | b2 | 0.13615 | 0.02874 |  | Data Grouping | by tip
5 | 6.3 | 916245643802300229 | e000231 | 12078.4395 | 11787.6508 | 13.35% | 87.7653 |  |  | #DIV/0! | 0% | 17.7 | 62148 | 19.4874 | 56638.0166 | desired ULOQ | 17.7 | mg/ml |  | b3 | -0.0030547 | 0.0014044
|  | 916245643802300230 | e000056 | 10088.8418 | [7 empty] | 17.7 | 44089.2514 | 14.76 | 56638.0166 | LLOD | Not Reported | mg/ml |  | b4 | #N/A | #N/A |  | Precision criteria for estimation of LLOQ | On
|  | 916245643802300231 | e000265 | 13195.671 | [11 empty] | LLOQ accuracy | 105.7764 | % |  | b5 | #N/A | #N/A
6 | 3.2 | 916245643802300232 | e000049 | 4506.8028 | 5378.655 | 22.92% | 40.0469 |  |  | #DIV/0! | 0% |  |  |  |  | LLOQ precision | 25.6974 | %
|  | 916245643802300233 | e000316 | 6250.5073 | [11 empty] | Average Residuals | 16.497 | % |  | 0 |  |  |  | Preferred models (if any) | Quadratic
|  | 916245643802300235 | e000169 | 9 | [11 empty] | Error in prediction: Best case | 12.2506 | %
7 | 0 | 916245643802300236 | e000161 | 151.9366 | 134.3088 | 23.2% | 1 |  |  | #DIV/0! | 0% |  |  |  |  | Error in prediction: Expected | 12.6167 | %
|  | 916245643802300237 | e000083 | 98.3315
|  | 916245643802300238 | e000188 | 152.6582 | [11 empty] | Signal Min | 4253.2433 | RLU | [5 empty] | fingerstickmultiplier | 1
[16 empty] | Signal Max | 63967.9448 | RLU |  |  |  | Quadratic |  | venousmultiplier | 1
[24 empty] | plasmamultiplier | 1
[16 empty] | Outliers detected based on within-standard variance and manual detection | Outliers |  |  |  | ULOQ accuracy | 96.744 | % | srerun_lo | 50
[17 empty] | Conc | RLU | 1 |  | ULOQ precision | 19.5212 | % | srerun_hi | 500000
[16 empty] | Statistically Detected ■ Manually Selected ■ | 0 | 98.3315 | 1
[17 empty] | 11.1 | 1 | 1
[17 empty] | 3.2 | 1 | 0
[17 empty] | 8 | 19548.4531 | 0
[17 empty] | 9.2 | 1 | 0
[17 empty] | 17.7 | 28157 | 0

## TSH Calibration
TSH Calibration
7-10-13 SZ
Female Health Demo Cartridges
Protocol: Female_Health svn
Calibrators: TSH Calibrators DOM 6-4-13 NN
Conclusion
1.
2. released status:
| TSH Calibration (includes re-run of cal# 2,4) : 06-27-13 AK + 07-09-13 SZ | [12 empty] | Data |  | Model Predictions |  | Model Type | 4PL | [6 empty] | Date | 2013-07-10
[9 empty] | Back Calculation |  |  |  | Conc | RLU | RLU → Conc | Conc → RLU | Model Equation | RLU = b1+(b2-b1)/(1+(Conc/b3)^b4 | [6 empty] | Name | SZ
| Calibrator # | Nominal Conc uIU/ml | Barcode | Edison # | Tip 3 | Mean | CV% | Modulation | Tip 3 | Mean | CV% | %Recovery | 0 | 164.5037 | OORL | 69.4849 | Calibration Equation | If RLU <= SignalMax & RLU >= SignalMin Then Conc = b3 * (((b2 - b1) / (RLU - b1)) - 1) ^ (1 / b4) End If | [6 empty] | Cartridge Lot | 2456485018
| 2* | 50 | 916245643802300349 | e000327 | 15242.6092 | 14041.3319 | 10.75% | 132.8733 | 53.0913 | 45.6892 | 20.41% | 91.38% | 0 | 9 | OORL | 69.4849 | [8 empty] | Analyte | TSH
|  |  | 916245643802300350 | e000198 | 12346.8792 |  |  |  | 32.2607 |  |  |  | 0 | 46.8455 | OORL | 69.4849 | [8 empty] | Units | uIU/ml
|  |  | 916245643802300351 | e000312 | 14534.5074 |  |  |  | 46.7427 |  |  |  | 0.1 | 91.5934 | OORL | 125.7058 | [8 empty] | Cartridge Id | 124
|  |  | 901245647105000001 | e000198 | 14985.5268 |  |  |  | 50.6621 |  |  |  | 0.1 | 106.4864 | OORL | 125.7058 | [8 empty] | Dilution | 10
|  |  | 901245647105000002 | e000331 | 17072.2112 |  |  |  | OORH |  |  |  | 0.1 | 142.153 | OORL | 125.7058 | [8 empty] | Reagent Lot | 2456485018
|  |  | 901245647105000003 | e000312 | 37044 | [7 empty] | 0.25 | 315 | 0.40218 | 218.0787
| 4* | 10 | 916245643802300352 | e000343 | 5369.3489 | 5966.4313 | 14.15% | 56.4604 | 9.1206 | 10.8052 | 13.57% | 108.1% | 0.5 | 351.3963 | 0.45861 | 378.2096 | [8 empty] | desired LLOQ | 0.5
|  |  | 916245643802300353 | e000060 | 6563.5136 |  |  |  | 11.7911 |  |  |  | 0.5 | 456.0833 | 0.61957 | 378.2096 | [8 empty] | desired ULOQ | 50
|  |  | 916245643802300354 | e000180 | 2227 | [7 empty] | 1 | 563.7551 | 0.78384 | 706.0724 | [8 empty] | max LLOQ | 0.5
|  |  | 901245647105000004 | e000180 | 2626.038 | [7 empty] | 1 | 764.4332 | 1.0885 | 706.0724 | LLOQ | 0.5 | uIU/ml |  | Model Parameters |  | SE |  | min ULOQ | 50
|  |  | 901245647105000005 | e000102 | 2637.2086 | [7 empty] | 1 | 614.2391 | 0.8606 | 706.0724 | ULOQ | 50 | uIU/ml |  | b1 | 69.4849 | 31.0393
|  |  | 901245647105000006 | e000163 | 6442 |  |  |  | 11.5038 |  |  |  | 5 | 2970.8873 | 4.5945 | 3206.913 | desired LLOQ | 0.5 | uIU/ml |  | b2 | 22953.332 | 3802.5988 |  | Data Grouping | by tip
| 5 | 5 | 901245647105000007 | e000060 | 2970.8873 | 2929.8722 | 1.98% | 27.7254 | 4.5945 | 4.5249 | 2.176% | 90.5% | 5 | 2888.8571 | 4.4553 | 3206.913 | desired ULOQ | 50 | uIU/ml |  | b3 | 28.1205 | 9.509
|  |  | 901245647105000008 | e000139 | 1610 | [7 empty] | 10 | 5369.3489 | 9.1206 | 5779.0633 | LLOD | Not Reported | uIU/ml |  | b4 | -1.0651 | 0.087077 |  | Precision criteria for estimation of LLOQ | On
|  |  | 901245647105000009 | e000093 | 2888.8571 |  |  |  | 4.4553 |  |  |  | 10 | 6563.5136 | 11.7911 | 5779.0633 | LLOQ accuracy | 107.8181 | % |  | b5 | #N/A | #N/A
| 7 | 1 | 901245647105000010 | e000290 | 563.7551 | 647.4758 | 16.12% | 6.1271 | 0.78384 | 0.91099 | 17.39% | 91.1% | 10 | 6442 | 11.5038 | 5779.0633 | LLOQ precision | 21.1135 | %
|  |  | 901245647105000011 | e000259 | 764.4332 |  |  |  | 1.0885 |  |  |  | 50 | 15242.6092 | 53.0913 | 14912.6021 | Average Residuals | 10.6697 | % |  | 0 |  |  |  | Preferred models (if any) | 4PL
|  |  | 901245647105000012 | e000033 | 614.2391 |  |  |  | 0.8606 |  |  |  | 50 | 12346.8792 | 32.2607 | 14912.6021 | Error in prediction: Best case | 15.9482 | %
| 8 | 0.5 | 901245647105000013 | e000133 | 221.3218 | 342.9338 | 34.3% | 3.2452 |  | 0.53909 | 21.11% | 107.8% | 50 | 14534.5074 | 46.7427 | 14912.6021 | Error in prediction: Expected | 18.8567 | %
|  |  | 901245647105000014 | e000182 | 351.3963 |  |  |  | 0.45861 |  |  |  | 50 | 14985.5268 | 50.6621 | 14912.6021
|  |  | 901245647105000015 | e000217 | 456.0833 |  |  |  | 0.61957 |  |  |  | 50 | 17072.2112 | OORH | 14912.6021 | Signal Min | 313.5971 | RLU | [5 empty] | fingerstickmultiplier | 1.6
| 9 | 0.25 | 901245647105000016 | e000061 | 103 | 315 |  | 2.9808 | OORL | [7 empty] | Signal Max | 15670.8319 | RLU | [5 empty] | venousmultiplier | 1.6
|  |  | 901245647105000017 | e000284 | 9 | [17 empty] | 4PL |  | plasmamultiplier | 1
|  |  | 901245647105000018 | e000328 | 315 | [11 empty] | Outliers detected based on within-standard variance and manual detection | Outliers |  |  |  | ULOQ accuracy | 103.5776 | % | srerun_lo | 50
| 10 | 0.1 | 901245647105000019 | e000229 | 91.5934 | 113.4109 | 22.91% | 1.0732 | OORL | [8 empty] | Conc | RLU | 1 |  | ULOQ precision | 30.6078 | % | srerun_hi | 1000000
|  |  | 901245647105000020 | e000221 | 106.4864 | [11 empty] | Statistically Detected ■ Manually Selected ■ | 50 | 37044 | 1
|  |  | 901245647105000021 | e000326 | 142.153 | [12 empty] | 0.25 | 9 | 0
| 12 | 0 | 901245647105000022 | e000315 | 164.5037 | 105.6746 | 78.73% |  | OORL | [8 empty] | 0.25 | 103 | 0
|  |  | 901245647105000023 | e000165 | 9 | [12 empty] | 0.5 | 221.3218 | 0
|  |  | 901245647105000024 | e000111 | 46.8455 | [12 empty] | 5 | 1610 | 0
[19 empty] | 10 | 2227 | 0
[19 empty] | 10 | 2637.2086 | 0
[19 empty] | 10 | 2626.038 | 0

## Vit D Calibration Serum
Vitamin D Calibration
7-8-13 SZ
Male Health Demo Cartridges
Protocol: Male_Health svn
Calibrators: Vitamin D Calibrators in Serum
Conclusion
1. The CVs and modulation are much better than the last calibration in 6-29-13
2. released status:
Vitamin D Calibration | [11 empty] | Data |  | Model Predictions |  | Model Type | 4PL | [6 empty] | Date | 2013-07-08
[8 empty] | Back Calculation |  |  |  | Conc | RLU | RLU → Conc | Conc → RLU | Model Equation | RLU = b1+(b2-b1)/(1+(Conc/b3)^b4 | [6 empty] | Name | SZ
Calibrator # | Re-assigned Conc by CLIA ng/ml | Barcode | Edison # | Tip 5 | Mean | CV | Modulation | Tip 5 | Mean | CV% | %Recovery | 5.5 | 5361.6984 | OORH | 5492.6996 | Calibration Equation | If RLU <= SignalMax & RLU >= SignalMin Then Conc = b3 * (((b2 - b1) / (RLU - b1)) - 1) ^ (1 / b4) End If | [6 empty] | Cartridge Lot | 1234
1 | 179.9 | 916245643802300325 | e000268 | 670.2522 | 652.8752 | 9.225% | 9.5015 | 161.6645 | 171.107 | 15.45% | 95.11% | 5.5 | 5999.2494 | OORH | 5492.6996 | [8 empty] | Analyte | Vitamin D (Serum)
|  | 916245643802300326 | e000339 | 702.5038 |  |  |  | 150.6806 |  |  |  | 11.1 | 4281.8092 | 12.4174 | 4485.7786 | [8 empty] | Units | ng/ml
|  | 916245643802300327 | e000280 | 585.8696 |  |  |  | 200.9758 |  |  |  | 11.1 | 4038.1657 | 14.113 | 4485.7786 | [8 empty] | Cartridge Id | 118
2 | 118.4 | 916245643802300328 | e000071 | 517 | 566.2122 | 12.29% | 10.9557 | 253.7386 | 219.3706 | 22.16% | 185.3% | 19.6 | 2944.7772 | 24.3211 | 3380.3855 | [8 empty] | Dilution | 50
|  | 916245643802300329 | e000187 | 903 | [7 empty] | 19.6 | 3638.7853 | 17.2559 | 3380.3855 | [8 empty] | Reagent Lot | 1234
|  | 916245643802300330 | e000115 | 615.4243 |  |  |  | 185.0025 |  |  |  | 19.6 | 3799.3256 | 15.9307 | 3380.3855
3 | 90.2 | 916245643802300331 | e000095 | 1014.3836 | 996.0634 | 7.948% | 6.2278 | 92.2602 | 95.0332 | 10.29% | 105.4% | 33.7 | 1869.4004 | 44.2859 | 2324.5575 | [8 empty] | desired LLOQ | 11.1
|  | 916245643802300332 | e000230 | 1064.462 |  |  |  | 86.9378 |  |  |  | 33.7 | 2320.4096 | 33.7783 | 2324.5575 | [8 empty] | desired ULOQ | 179.9
|  | 916245643802300333 | e000140 | 909.3447 |  |  |  | 105.9016 |  |  |  | 33.7 | 2743.9776 | 26.9333 | 2324.5575 | [8 empty] | max LLOQ | 11.1
4 | 61.9 | 916245643802300334 | e000056 | 1461.6657 | 1498.3969 | 18.28% | 4.1399 | 59.4555 | 59.3788 | 21.26% | 95.93% | 61.9 | 1461.6657 | 59.4555 | 1412.8425 | LLOQ | 11.1 | ng/ml |  | Model Parameters |  | SE |  | min ULOQ | 179.9
|  | 916245643802300335 | e000265 | 1788.788 |  |  |  | 46.7173 |  |  |  | 61.9 | 1788.788 | 46.7173 | 1412.8425 | ULOQ | 179.9 | ng/ml |  | b1 | 317.825 | 1118.5163
|  | 916245643802300336 | e000355 | 1244.7371 |  |  |  | 71.9635 |  |  |  | 61.9 | 1244.7371 | 71.9635 | 1412.8425 | desired LLOQ | 11.1 | ng/ml |  | b2 | 6460.3417 | 202.3222 |  | Data Grouping | by tip
5 | 33.7 | 916245643802300337 | e000049 | 1869.4004 | 2311.2625 | 18.92% | 2.6839 | 44.2859 | 34.9992 | 24.97% | 103.9% | 90.2 | 1014.3836 | 92.2602 | 1033.0625 | desired ULOQ | 179.9 | ng/ml |  | b3 | 19.5163 | 5.7468
|  | 916245643802300338 | e000316 | 2320.4096 |  |  |  | 33.7783 |  |  |  | 90.2 | 1064.462 | 86.9378 | 1033.0625 | LLOD | Not Reported | ng/ml |  | b4 | 1.3239 | 0.30516 |  | Precision criteria for estimation of LLOQ | On
|  | 916245643802300339 | e000231 | 2743.9776 |  |  |  | 26.9333 |  |  |  | 90.2 | 909.3447 | 105.9016 | 1033.0625 | LLOQ accuracy | 119.5061 | % |  | b5 | #N/A | #N/A
6 | 19.6 | 916245643802300340 | e000163 | 2944.7772 | 3460.9627 | 13.12% | 1.7924 | 24.3211 | 19.1692 | 23.53% | 97.8% | 118.4 | 903 | 106.86 | 834.9657 | LLOQ precision | 9.0382 | %
|  | 916245643802300341 | e000327 | 3638.7853 |  |  |  | 17.2559 |  |  |  | 118.4 | 615.4243 | 185.0025 | 834.9657 | Average Residuals | 11.1778 | % |  | 0 |  |  |  | Preferred models (if any) | 4PL
|  | 916245643802300342 | e000198 | 3799.3256 |  |  |  | 15.9307 |  |  |  | 179.9 | 670.2522 | 161.6645 | 626.0866 | Error in prediction: Best case | 16.7178 | %
7 | 11.1 | 916245643802300343 | e000343 | 4281.8092 | 4485.1083 | 12.85% | 1.3831 | 12.4174 | 13.2652 | 9.038% | 119.5% | 179.9 | 702.5038 | 150.6806 | 626.0866 | Error in prediction: Expected | 19.9833 | %
|  | 916245643802300344 | e000180 | 4038.1657 |  |  |  | 14.113 |  |  |  | 179.9 | 585.8696 | 200.9758 | 626.0866
|  | 916245643802300345 | e000197 | 5135.3499 | [11 empty] | Signal Min | 576.2054 | RLU | [5 empty] | fingerstickmultiplier | 1
8 | 5.5 | 916245643802300346 | e000312 | 7248.8825 | 6203.2768 | 15.48% | 1 |  | 4.55 | 50.42% | 82.73% |  |  |  |  | Signal Max | 4859.1788 | RLU | [5 empty] | venousmultiplier | 1
|  | 916245643802300347 | e000060 | 5361.6984 |  |  |  | 6.172 | [13 empty] | 4PL |  | plasmamultiplier | 1
|  | 916245643802300348 | e000217 | 5999.2494 |  |  |  | 2.9279 | [7 empty] | Outliers detected based on within-standard variance and manual detection | Outliers |  |  |  | ULOQ accuracy | 95.1123 | % | srerun_lo | 50
[17 empty] | Conc | RLU | 1 |  | ULOQ precision | 15.4545 | % | srerun_hi | 10000
[16 empty] | Statistically Detected ■ Manually Selected ■ | 5.5 | 7248.8825 | 0
[17 empty] | 11.1 | 5135.3499 | 0
[17 empty] | 118.4 | 517.0486 | 0

## tPSA_MaleHealth_Calibration
tPSA Male Health Calibration 07-09-2013 AK (Re-Run of Calibrator 1 and 6) | [13 empty] | Historical Data: 04/02/13 NN
Back Calculation
Calibrator # | Concentration (ng/mL) |  | Barcode | Edison # | Tip 4 | Mean | CV% | Modulation | Tip 4 | Mean | CV% | %Recovery |  | Male Health Cartridge Calibration
1 | 160 |  | 901245647101200217 | e000165 | 583851.7799 | 465721.619 | 22.28% | 270.8924 | #NUM! | 156.2165 | 21.25% | 97.64% |  | 04/02/13 NN
|  |  | 901245647101200218 | e000061 | 389297.736 |  |  |  | 132.7447 |  |  |  |  | Calibrators | Calibrators made LW
|  |  | 901245647101200219 | e000284 | 424015.341 |  |  |  | 179.6884 |  |  |  |  | Cartridge Lot# | ; Cartridges provided by EC 04/02/13
2 | 40 |  | 901245647101200196 | e000115 | 195223.0432 | 236320.5492 | 19.07% | 137.4586 | 30.5968 | 43.5471 | 34.62% | 108.9% |  | Protocol: | Male_Health svn 6907 | [5 empty] | Back-Calculation
|  |  | 901245647101200197 | e000187 | 229229.068 |  |  |  | 39.9503 |  |  |  |  | CLIA Assigned Values tPSA Conc ng/mL | Barcode# | Edison# | RLU | AVG RLU | CV | Modulation | Conc | Mean | % CV | % Recovery
|  |  | 901245647101200198 | e000230 | 284509.5363 |  |  |  | 60.0943 |  |  |  |  | 160 | 118245627502300055 | e000169 | 1121040.3066 | 952296.6874 | 20.7% | 302.5606 | OORH | #DIV/0! | #DIV/0! | #DIV/0!
3 | 10 |  | 901245647101200199 | e000095 | 77783.0621 | 81259.9104 | 3.807% | 47.2658 | 8.7202 | 9.2052 | 4.691% | 92.05% |  |  | 118245627502300056 | e000022 | 735566.1616
|  |  | 901245647101200200 | e000355 | 83707.9418 |  |  |  | 9.5479 | [5 empty] | 118245627502300057 | e000316 | 1000283.5939 |  |  |  | OORH
|  |  | 901245647101200201 | e000265 | 82288.7272 |  |  |  | 9.3475 |  |  |  |  | 40 | 118245627502300058 | e000165 | 555312.492 | 502881.4998 | 9.061% | 159.7739 | OORH | #DIV/0! | #DIV/0! | #DIV/0!
4 | 2.5 |  | 901245647101200202 | e000231 | 29985.3837 | 28597.651 | 6.079% | 16.6342 | 2.834 | 2.6837 | 6.998% | 107.3% |  |  | 118245627502300059 | e000188 | 480468.2407 |  |  |  | OORH
|  |  | 901245647101200203 | e000056 | 26647.8195 |  |  |  | 2.4732 | [5 empty] | 118245627502300060 | e000334 | 472863.7667
|  |  | 901245647101200204 | e000316 | 29159.7497 |  |  |  | 2.7441 |  |  |  |  | 10 | 118245627502300061 | e000108 | 134362.7097 | 137593.3116 | 7.959% | 43.7157 | OORL | #DIV/0! | #DIV/0! | #DIV/0!
5 | 0.625 |  | 901245647101200205 | e000188 | 9108.3297 | 8296.6487 | 8.584% | 4.8258 | 0.69388 | 0.61726 | 10.89% | 98.76% |  |  | 118245627502300062 | e000099 | 128621.1827 |  |  |  | OORL
|  |  | 901245647101200206 | e000083 | 8005.1171 |  |  |  | 0.58968 | [5 empty] | 118245627502300063 | e000071 | 149796.0423 |  |  |  | OORL
|  |  | 901245647101200207 | e000100 | 7776.4994 |  |  |  | 0.56823 |  |  |  |  | 2.5 | 118245627502300064 | e000201 | 46505.2563 | 35332.7777 | 28% | 11.2258 | OORL | #DIV/0! | #DIV/0! | #DIV/0!
6 | 0.156 |  | 901245647101200220 | e000221 | 3400.5916 | 2274.262 | 85.21% | 1.3229 | 0.17029 | 0.16963 | 0.5454% | 108.7% |  |  | 118245627502300065 | e000309 | 31808.4171 |  |  |  | OORL
|  |  | 901245647101200221 | e000229 | 3385.6121 |  |  |  | 0.16898 | [5 empty] | 118245627502300066 | e000230 | 27684.6596 |  |  |  | OORL
|  |  | 901245647101200223 | e000328 | 36.5823 |  |  |  | #NUM! |  |  |  |  | 0.625 | 118245627502300067 | e000095 | 12316.485 | 10722.1792 | 16.05% | 3.4066 | OORL | #DIV/0! | #DIV/0! | #DIV/0!
7 | 0.039 |  | 901245647101200211 | e000164 | 1757.2581 | 1757.2581 |  | 1.0221 | 0.031127 | 0.031127 |  | 79.81% |  |  | 118245627502300068 | e000242 | 10952.273 |  |  |  | OORL
|  |  | 901245647101200212 | e000114 | 2 |  |  |  | #NUM! | [5 empty] | 118245627502300069 | e000054 | 8897.7796 |  |  |  | OORL
[14 empty] | 0.15625 | 118245627502300070 | e000199 | 6402.9115 | 6843.4792 | 6.267% | 2.1743 | OORL | #DIV/0! | #DIV/0! | #DIV/0!
8 | 0 |  | 901245647101200214 | e000237 | 1735.5659 | 1719.2126 | 8.153% |  | 0.029385 | 0.028147 | 39.75% |  |  |  | 118245627502300071 | e000280 | 6867.9222 |  |  |  | OORL
|  |  | 901245647101200215 | e000341 | 1850.4888 |  |  |  | 0.038665 | [5 empty] | 118245627502300072 | e000339 | 7259.6039 |  |  |  | OORL
|  |  | 901245647101200216 | e000041 | 1571.5832 |  |  |  | 0.016391 |  |  |  |  | 0.039062 | 118245627502300073 | e000264 | 2191.9108 | 2814.411 | 19.19% | 0.89419 | OORL | #DIV/0! | #DIV/0! | #DIV/0!
[15 empty] | 118245627502300074 | e000290 | 3093.5004 |  |  |  | OORL
[15 empty] | 118245627502300075 | e000164 | 3157.8219 |  |  |  | OORL
[14 empty] | 0 | 118245627502300076 | e000085 | 3578.1758 | 3147.458 | 14.98% |  | OORL | OORL
[15 empty] | 118245627502300077 | e000144 | 3220.4437 |  |  |  | OORL
[15 empty] | 118245627502300078 | e000140 | 2643.7545 |  |  |  | OORL
Data |  | Model Predictions |  |  |  | Model Type | 4PL | [6 empty] | Date | 2013-07-09
Conc | RLU | RLU → Conc | Conc → RLU |  |  | Model Equation | RLU = b1+(b2-b1)/(1+(Conc/b3)^b4 | [6 empty] | Name | AK
0 | 1757.2581 | OORL | 1351.6088 |  |  | Calibration Equation | If RLU <= SignalMax & RLU >= SignalMin Then Conc = b3 * (((b2 - b1) / (RLU - b1)) - 1) ^ (1 / b4) End If | [6 empty] | Cartridge Lot | 123
0 | 2 | OORL | 1351.6088 | [10 empty] | Analyte | tPSA
0 | 1735.5659 | OORL | 1351.6088 | [10 empty] | Units | ng/mL
0 | 1850.4888 | OORL | 1351.6088 | [10 empty] | Cartridge Id | 118
0 | 1571.5832 | OORL | 1351.6088 | [10 empty] | Dilution | 10
0.2 | 3400.5916 | 0.17029 | 3739.2296 | [10 empty] | Reagent Lot | 123
0.2 | 3385.6121 | 0.16898 | 3739.2296
0.6 | 9108.3297 | 0.69388 | 8114.9536 | [10 empty] | desired LLOQ | 0.04
0.6 | 8005.1171 | 0.58968 | 8114.9536 | [10 empty] | desired ULOQ | 160
0.6 | 7776.4994 | 0.56823 | 8114.9536 | [10 empty] | max LLOQ | 0.04
2.5 | 29985.3837 | 2.834 | 26898.0904 |  |  | LLOQ | 0.2 | ng/mL |  | Model Parameters |  | SE |  | min ULOQ | 160
2.5 | 26647.8195 | 2.4732 | 26898.0904 |  |  | ULOQ | 160 | ng/mL |  | b1 | 1351.6088 | 250.5364
2.5 | 29159.7497 | 2.7441 | 26898.0904 |  |  | desired LLOQ | 0.2 | ng/mL |  | b2 | 577950.7631 | 82784.1213 |  | Data Grouping | by tip
10 | 77783.0621 | 8.7202 | 86875.287 |  |  | desired ULOQ | 160 | ng/mL |  | b3 | 62.3821 | 18.964
10 | 83707.9418 | 9.5479 | 86875.287 |  |  | LLOD | Not Reported | ng/mL |  | b4 | -0.95472 | 0.053965 |  | Precision criteria for estimation of LLOQ | On
10 | 82288.7272 | 9.3475 | 86875.287 |  |  | LLOQ accuracy | 84.817 | % |  | b5 | #N/A | #N/A
40 | 195223.0432 | 30.5968 | 229392.7917 |  |  | LLOQ precision | 0.54538 | %
40 | 229229.068 | 39.9503 | 229392.7917 |  |  | Average Residuals | 7.6667 | % |  | 0 |  |  |  | Preferred models (if any) | 4PL
40 | 284509.5363 | 60.0943 | 229392.7917 |  |  | Error in prediction: Best case | 12.7671 | %
160 | 389297.736 | 132.7447 | 411195.3611 |  |  | Error in prediction: Expected | 16.6256 | %
160 | 424015.341 | 179.6884 | 411195.3611
[6 empty] | Signal Min | 3282.636 | RLU | [5 empty] | fingerstickmultiplier | 1.6
[6 empty] | Signal Max | 426556.2841 | RLU | [5 empty] | venousmultiplier | 1.6
[12 empty] | 4PL |  | plasmamultiplier | 1
[6 empty] | Outliers detected based on within-standard variance and manual detection | Outliers |  |  |  | ULOQ accuracy | 97.6353 | % | srerun_lo | 50
[7 empty] | Conc | RLU | 1 |  | ULOQ precision | 21.2489 | % | srerun_hi | 5000000
[6 empty] | Statistically Detected ■ Manually Selected ■ | 0 | 1 | 1
[7 empty] | 0.2 | 36.5823 | 1
[7 empty] | 160 | 583851.7799 | 0