## Problems Summary
Total cartridges run (approx): |  |  |  | 900
Total Failures Due to Observed Problems |  |  |  | 49 | Does not include control failures
% of Cartridges Failed due to Observed Problems |  |  |  | 5.4444 | %
Number of cartridges with control failures |  |  |  | 100
% of Cartridges Failed due to Control failures |  |  |  | 11.1111 | %
% Total failure rate |  |  |  | 16.5556 | %
Unexected low counts likely due to protocol sample prep error (not explained by observed cartridge problems, close duplicate tips)
Total | % of Failures | Barcode | Reader | Tip 1 | Tip 2 | Tip 3 | Tip 4 | Tip 5 | Tip 6 | Observation
3 | 6.1224 | 060245551800600962 | e000071 | 5611.0263 | 8203.6734 | 35106.243 | 40743.1834 | 4306.6817 | 4544.8449 | unexpected 2x low sample counts (tip 1 and 2)
|  | 060245551800600605 | e000323 | 21979.5547 | 26496.3897 | 34575.1098 | 35824.8198 | 3602.9581 | 3968.3881 | unexpected 2x low sample counts (tip 1 and 2)
|  | 060245551800600299 | e000278 | 2716.3679 | 3023.0978 | 36565.8213 | 37540.9277 | 2325.9235 | 2646.0477 | unexpected 2x low high control counts (tip 5 and 6)
Unexpected high counts likely due to protocol sample prep error (not explained by observed cartridge problems, close duplicate tips)
Total | % of Failures | Barcode | Reader | Tip 1 | Tip 2 | Tip 3 | Tip 4 | Tip 5 | Tip 6 | Observation
4 | 8.1633 | 060245551800600915 | e000233 | 50539.1874 | 47679.8154 | 100187.8679 | 102173.6331 | 4444.7341 | 4613.0854 | unexpected 2x high counts in low control (tip 3 and 4)
|  | 060245551800600952 | e000303 | 243909.325 | 198278.0212 | 50123.4787 | 54494.0394 | 5541.4113 | 5482.701 | unexpected 4x high sample counts (tip 1 and 2)
|  | 060245551800600634 | e000188 | 73771.8592 | 88261.1834 | 65737.0878 | 64694.8863 | 4866.0826 | 4632.5569 | unexpected 1.5x high counts in low control (tip 3 and 4)
|  | 060245551800600669 | e000127 | 73404.602 | 71478.5413 | 62564.304 | 60050.233 | 4532.9067 | 5175.798 | unexpected 1.5x high counts in low control (tip 3 and 4)
Reagent Aspiration Problems
Total | % of Failures | Barcode | Reader | Tip 1 | Tip 2 | Tip 3 | Tip 4 | Tip 5 | Tip 6 | Observation
4 | 8.1633 | 060245551800600912 | e000134 | 29540.3857 | 47410.0422 | 55449.3417 | 51967.9536 | 5157.0291 | 5756.1468 | Low counts on tip 1 - incomplete aspiration of sample/conjugate mixture as evidenced by excess volume left in sample tube after run
|  | 060245551800600533 | e000143 | 65092.2187 | 83990.2675 | 51034.5459 | 44746.5555 | 5316.0985 | 4729.9646 | Low counts on tip 1 - incomplete aspiration of sample/conjugate mixture as evidenced by excess volume left in sample tube after run
|  | 060245551800600461 | e000290 | 69038.312 | 73565.5352 | 46259.0074 | 44507.1938 | 2231.2918 | 3650.9072 | bubble in tip 5 position
|  | 060245551800600108 | e000248 | 44895.7987 | 48155.8724 | 49413.9331 | 46973.9997 | 3422.053 | 403.9879 | Bubble in bottom of sample well tip 6 (airlock) - improper aspiration
Blood tip(s) fell off or was shucked
Total | % of Failures | Barcode | Reader | Tip 1 | Tip 2 | Tip 3 | Tip 4 | Tip 5 | Tip 6 | Observation
10 | 20.4082 | 060245551800600502 | e000341 | 53.2302 | 66.5378 | 40948.8112 | 44549.2429 | 4190.3998 | 4092.0718 | Blood tip #2 fell off or was shucked 2 blood tips present when cartridge loaded, blood tip 1 completed dilutions and put into F1, BT2 picked up but fell off/was shucked before diluent is to be picked up (foil on diluent for sample not pierced)
|  | 060245551800600537 | e000203 | 73.2328 | 68.3096 | 46956.3969 | 43393.8352 | 4410.5855 | 4232.7344 | Blood tip #2 fell off or was shucked 2 blood tips present when cartridge loaded, blood tip 1 completed dilutions and put into F1, BT2 picked up but fell off/was shucked before diluent is to be picked up (foil on diluent for sample not pierced)
|  | 060245551800600590 | e000134 | 76.5947 | 86.0379 | 50577.7067 | 52556.5789 | 3854.9186 | 4354.3582 | Blood tip #2 fell off or was shucked 2 blood tips present when cartridge loaded, blood tip 1 completed dilutions and put into F1, BT2 picked up but fell off/was shucked before diluent is to be picked up (foil on diluent for sample not pierced)
|  | 060245551800600584 | e000310 | 59.8606 | 61.4569 | 66.2457 | 55.8699 | 55.0718 | 67.842 | Both blood tips present when cartridge loaded, both BTs in bottom of reader after run and no dilutions done
|  | 060245551800600678 | e000132 | 61171.6461 | 65942.3673 | 102.8276 | 106.4146 | 104.0232 | 101.6319 | Both blood tips present when cartridge loaded, tip #1 in bottom of reader after run and no dilutions done for controls
|  | 060245551800600720 | e000290 | 59.7302 | 58.0938 | 41405.3136 | 46970.0425 | 2991.4202 | 3907.0108 | Blood tip #2 fell off or was shucked 2 blood tips present when cartridge loaded, blood tip 1 completed dilutions and put into F1, BT2 picked up but fell off/was shucked before diluent is to be picked up (foil on diluent for sample not pierced)
|  | 060245551800600675 | e000140 | 72.3856 | 66.7305 | 47918.1682 | 43917.7301 | 4461.8966 | 4940.3205 | Blood tip #2 fell off or was shucked 2 blood tips present when cartridge loaded, blood tip 1 completed dilutions and put into F1, BT2 picked up but fell off/was shucked before diluent is to be picked up (foil on diluent for sample not pierced)
|  | 060245551800600719 | e000130 | 65.4649 | 65.4649 | 37985.7457 | 40241.7662 | 3409.2096 | 3532.0821 | Blood tip #2 fell off or was shucked 2 blood tips present when cartridge loaded, blood tip 1 completed dilutions and put into F1, BT2 picked up but fell off/was shucked before diluent is to be picked up (foil on diluent for sample not pierced)
|  | 060245551800600312 | e000275 | 56.0158 | 52.3626 | 42449.049 | 43211.9604 | 2547.5031 | 3201.4272 | Blood tip #2 fell off or was shucked 2 blood tips present when cartridge loaded, blood tip 1 completed dilutions and put into F1, BT2 picked up but fell off/was shucked before diluent is to be picked up (foil on diluent for sample not pierced)
|  | 060245551800600331 | e000327 | 59.474 | 62.032 | 40039.4279 | 42090.3214 | 2995.4429 | 3657.9704 | Blood tip #2 fell off or was shucked 2 blood tips present when cartridge loaded, blood tip 1 completed dilutions and put into F1, BT2 picked up but fell off/was shucked before diluent is to be picked up (foil on diluent for sample not pierced)
Tip Pickup Problems (not obviously related to cartridge problem)
Total | % of Failures | Barcode | Reader | Tip 1 | Tip 2 | Tip 3 | Tip 4 | Tip 5 | Tip 6 | Observation
3 | 6.1224 | 060245551800600960 | e000056 | 1468.5603 | 1589.7091 | 37333.541 | 6.8946 | 4203.7662 | 4125.9551 | tip 4 not picked up
|  | 060245551800600910 | e000275 | 42073.3775 | 38924.313 | 5.4798 | 46003.6195 | 3919.2824 | 3653.2071 | tip 3 not picked up
|  | 060245551800600052 | e000209 | 83413.8427 | 85095.2052 | 42907.8347 | 40046.0837 | 3396.235 | 0.83775 | Tip 6 not picked up
Cartridge problems
Total | % of Failures | Barcode | Reader | Tip 1 | Tip 2 | Tip 3 | Tip 4 | Tip 5 | Tip 6 | Observation
19 | 38.7755 | 060245551800600957 | e000098 | 105083.8786 | 83343.0743 | 22427.9334 | 49388.5254 | 2793.3942 | 5759.3186 | Substrate tube #3&5 low volume
|  | 060245551800600958 | e000176 | 29350.5536 | 86035.7168 | 48194.1358 | 54517.6444 | 4736.7106 | 5415.9549 | Substrate tube #1 low volume
|  | 060245551800600981 |  | 0 | [5 empty] | Barcode partially rubbed off or printer problem - would not scan
|  | 060245551800600631 | e000132 | 105.2189 | 96.8492 | 101.6319 | 94.4579 | 4475.3909 | 3958.8616 | Conjugate in D7 (for low control tips 3-4) and F5 for sample tips 1-2) evaporated to very low level (< 20uL left)
|  | 060245551800600474 | e000310 | 62.255 | 60.6587 | 44692.7265 | 43920.1256 | 4319.5413 | 4805.6094 | F5 conjugate for tips 1-2 empty (evaporated?)
|  | 060245551800600152 | e000272 | 48400.5929 | 48683.0595 | 2.4246 | 1.8185 | 4494.0074 | 1.8185 | No substrate left in 3, 4 and 6
|  | 060245551800600155 | e000210 | 64.0277 | 78.3397 | 41987.8421 | 35179.8172 | 3201.3837 | 3148.655 | F5 conjugate for tips 1-2 empty (evaporated?)
|  | 060245551800600470 | e000140 | 66.7305 | 71.2546 | 43554.6708 | 43923.3852 | 3395.3393 | 4123.7199 | F5 conjugate for tips 1-2 empty (evaporated?)
|  | 060245551800600103 | e000188 | 107.2733 | 100.6719 | 53497.2025 | 44234.5643 | 3789.2236 | 3492.9842 | Diluent tube in F5 conjugate position (Black-blue color code instead of red)
|  | 060245551800600260 | e000209 | 54.4537 | 55.2914 | 49629.096 | 38280.9463 | 4187.9079 | 3698.6624 | F5 conjugate for tips 1-2 empty (evaporated?)
|  | 060245551800600262 | e000240 | 48.0027 | 49.7805 | 42842.6677 | 44340.2321 | 4044.0757 | 3815.3223 | F5 conjugate empty (evaporated?)
|  | 060245551800600224 | e000052 | 81.4852 | 63.2583 | 43388.7342 | 45434.4421 | 3462.0496 | 2998.8705 | F5 conjugate empty (evaporated?)
|  | 060245551800600107 | e000131 | 1412.1614 | 1339.248 | 42521.3924 | 44325.9996 | 3453.7373 | 3015.1845 | Conjugate in F5 position very low (< 50uL, should be ~90uL after run)
|  | 060245551800600789 | e000275 | 63324.6922 | 65880.1106 | 4.8709 | 42012.4907 | 3860.8311 | 4020.3544 | Tip 3 no substrate
|  | 060245551800600825 | e000289 | 70870.9367 | 72746.6025 | 57993.8007 | 22.7353 | 4250.6972 | 21.1114 | Tip 4 and 6 no substrate
|  | 060245551800600740 | e000277 | 1577.9337 | 1523.0427 | 51672.2058 | 47526.1045 | 4389.0856 | 253.9624 | Low substrate in tube #6
|  | 060245551800600175 | e000209 | 22302.5572 | 18943.1832 | 0.83775 | 48674.062 | 4181.2059 | 4282.5735 | Tip 3 not picked up - jammed tight in cartridge
|  | 060245551800600212 | e000228 | 58908.5961 | 71930.4963 | 20.9608 | 19.2141 | 3977.3128 | 4449.8043 | Tips 3-4 did not pick up, checked after run they were jammed in really tight and had to be pried out
|  | 060245551800600033 | e000176 | 7.5787 | 70.1033 | 39313.7507 | 35662.6942 | 3492.8504 | 2825.9215 | F5 conjugate for tips 1-2 empty (evaporated?)
Pump Failure
Total | % of Failures | Barcode | Reader
2 | 4.0816 | 060245551800600500 | e000268 | 0 | [5 empty] | pump failure - no data
|  | 060245551800600200 |  | 0 | [5 empty] | pump failure, no data
Reader-Specific Problems
Total | % of Failures | Barcode | Reader | Tip 1 | Tip 2 | Tip 3 | Tip 4 | Tip 5 | Tip 6 | Observation
4 | 8.1633 | 060245551800600949 | e000277 | 54716.0956 | 42336.342 | 40327.3308 | 39099.9678 | 4230.9994 | 1540.6078 | Reader 277 not aspirating/loading tip 6 properly, pipette broken
|  | 060245551800600637 | e000277 | 2238.0896 | 2481.0739 | 57415.2697 | 57353.0599 | 4300.528 | 1040.7336 | Reader 277 not aspirating/loading tip 6 properly, pipette broken
|  | 060245551800600513 | e000277 | 6635.9578 | 10020.9037 | 45662.7375 | 42627.6303 | 4506.9183 | 608.9243 | Reader 277 not aspirating/loading tip 6 properly, pipette broken
|  | 060245551800600917 | e000277 | 2351.5311 | 2539.6243 | 47382.656 | 46137.7277 | 4304.1875 | 327.1504 | Reader 277 not aspirating/loading tip 6 properly
Control Failures
|  | 060245551800600845 | e000132
|  | 060245551800600887 | e000210
|  | 060245551800601002 | e000143
|  | 060245551800600915 | e000233
|  | 060245551800600937 | e000289
|  | 060245551800600973 | e000071
|  | 060245551800600928 | e000278
|  | 060245551800600991 | e000192
|  | 060245551800600948 | e000071
|  | 060245551800600636 | e000140
|  | 060245551800600838 | e000071
|  | 060245551800600509 | e000323
|  | 060245551800600513 | e000277
|  | 060245551800600625 | e000337
|  | 060245551800600709 | e000248
|  | 060245551800600590 | e000134
|  | 060245551800600867 | e000277
|  | 060245551800600556 | e000323
|  | 060245551800600572 | e000327
|  | 060245551800600584 | e000310
|  | 060245551800600545 | e000069
|  | 060245551800600065 | e000209
|  | 060245551800600085 | e000052
|  | 060245551800600074 | e000265
|  | 060245551800600078 | e000130
|  | 060245551800600065 | e000209
|  | 060245551800600085 | e000052
|  | 060245551800600074 | e000265
|  | 060245551800600078 | e000130
|  | 060245551800600664 | e000323
|  | 060245551800600631 | e000132
|  | 060245551800600721 | e000278
|  | 060245551800600728 | e000310
|  | 060245551800600529 | e000071
|  | 060245551800600653 | e000303
|  | 060245551800600678 | e000132
|  | 060245551800600608 | e000203
|  | 060245551800600634 | e000188
|  | 060245551800600881 | e000337
|  | 060245551800600669 | e000127
|  | 060245551800600698 | e000323
|  | 060245551800600671 | e000253
|  | 060245551800600215 | e000214
|  | 060245551800600337 | e000191
|  | 060245551800600482 | e000233
|  | 060245551800600301 | e000191
|  | 060245551800600144 | e000183
|  | 060245551800600315 | e000278
|  | 060245551800600147 | e000278
|  | 060245551800600299 | e000278
|  | 060245551800600141 | e000227
|  | 060245551800600113 | e000022
|  | 060245551800600124 | e000027
|  | 060245551800600199 | e000269
|  | 060245551800600211 | e000132
|  | 060245551800600452 | e000085
|  | 060245551800600114 | e000269
|  | 060245551800600128 | e000006
|  | 060245551800600200 | e000202
|  | 060245551800600197 | e000289
|  | 060245551800600152 | e000272
|  | 060245551800600275 | e000289
|  | 060245551800600441 | e000098
|  | 060245551800600122 | e000233
|  | 060245551800600155 | e000210
|  | 060245551800600091 | e000209
|  | 060245551800600255 | e000071
|  | 060245551800600445 | e000323
|  | 060245551800600204 | e000248
|  | 060245551800600213 | e000240
|  | 060245551800600224 | e000052
|  | 060245551800600431 | e000191
|  | 060245551800600116 | e000275
|  | 060245551800600459 | e000194
|  | 060245551800600237 | e000341
|  | 060245551800600111 | e000327
|  | 060245551800600432 | e000253
|  | 060245551800600447 | e000278
|  | 060245551800600107 | e000131
|  | 060245551800600258 | e000131
|  | 060245551800600240 | e000323
|  | 060245551800600096 | e000204
|  | 060245551800600101 | e000323
|  | 060245551800600444 | e000027
|  | 060245551800600282 | e000323
|  | 060245551800600212 | e000228
|  | 060245551800600092 | e000039
|  | 060245551800600095 | e000130
|  | 060245551800600434 | e000198
|  | 060245551800600236 | e000148
|  | 060245551800600099 | e000148
|  | 060245551800600264 | e000039
|  | 060245551800600265 | e000148
|  | 060245551800600283 | e000297
|  | 060245551800600738 | e000277
|  | 060245551800600826 | e000327
|  | 060245551800600223 | e000052
|  | 060245551800600188 | e000337
|  | 060245551800600233 | e000290
|  | 060245551800600234 | e000275