## Assay list
| ASSAY |  | Status | Other
1 | thyroid stim hormone
2 | glycosylated hemoglobin
3 | vitamin d, 25 hydroxy
4 | psa, total
5 | thyroxine, free
6 | chorionic gonadotropin
7 | c-reactive protein
8 | ferritin
9 | vitamin b-12
10 | hepatitis b surface ag, eia
11 | syphilis non-trep qual;
12 | hiv-1/hiv-2, single Antibody; HIV-1 and HIV-2, single assay
13 | lyme disease antibody Antibody; Borrelia burgdorferi (Lyme disease
14 | troponin T, quant
15 | thyroxine, total
16 | hepatitis c ab Hepatitis C antibody;
17 | folic acid serum, blood
18 | antinuclear antibodies
19 | gonadotropin (fsh)
20 | total testosterone (DEA license pending)
21 | rubella antibody
22 | estradiol
23 | thyroid (t3 or t4)/Thyroid hormone (T3 or T4) uptake or thyroid hormone binding ratio (THBR)
24 | creatine, mb fraction
25 | prolactin
26 | triiodothyronine (t3) Triiodothyronine T3; total (TT-3)
27 | gonadotropin (lh)
28 | strep a w/optic
29 | pregnancy , urine Urine pregnancy test, by visual color comparison methods
| Code | Status | Number of Assays
|  | Assay complete | 11
|  | In development | 15
|  | Dose response demonstrated in Urine matrix, clinical correlation in urine pending | 1
|  | Assays for which clinical correlation requires Biosafety labs | 5
|  | Development to start | 2

## Summary checklist
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29
| Experiment | Information desired | TSH | HbA1c* | 25-OH Vit D* | Total PSA | Free T4 | hCG | CRP | Ferritin | Vit B-12 | HbS Ag | Non- treponemal Syphilis Test | HIV-1 & 2 Ab | Lyme disease Ab | cTnT | T4 Total | Hep C Ab | Folic acid | ANA | FSH | Testosterone, total | Rubella Ab | Estradiol | T3, T4 uptake | CK-MB | Prolactin | T3, total | LH | Strep A | Pregnancy, urine
1 | MTP screen | Summary information of antibody screens, 4 point curves
2 | Conjugation Type: SH or NH2 | Compare SH vs. NH2 conjugations (For assays with limited antibodies to screen) | 1: N/A; 5: N/A; 11: N/A; 15: N/A; 19: N/A; 21: NA; 22: NA
3 | Cross reactivity | For 4 best Ab. Pairs:Test cross reactivity with related analytes, 6 point standard curves | 11: N/A; 21: NA
4 | Interference | For 4 best Ab. Pairs:Test interference with related analytes, 6 point standard curves | 5: N/A; 11: N/A; 13: N/A; 21: NA
5 | Epitope map | For 4 best Ab. Pairs | 15: N/A; 21: NA; 25: N/A
6 | Theranos screen | For 3 best Ab. Pairs: Summary information for 6 point curves in A.B, whole blood and plasma | bar: 4 to 8; 10 to 15; 18 to 19; 25; 27 to 29
7 | Training set | For 1 Ab. Pair correlation of N=15 clinical samples spanning the assay range range | bar: 3 to 6; 8; 10 to 15; 18 to 19; 27 to 28
8 | Calibrator comparison/screen | Run our controls on reference kit and vice versa, calibrate with WHO standards | bar: 3 to 15; 21 to 23; 29
Antibody Pair finalized; Order 3 mg of final pair | bar: 4 to 7; 11; 13; 25; 29
9 | Plasma Screen | Screen normal plasma to see ranges, Minimum N =10 | 29: N/A
10 | Whole Blood Screen | Screen normal blood to see ranges. Minimum N=10 | 15: N/A; 29: N/A
11 | Capture titration | Show table with selected level and higher/lower levels at 20, 10, 5, 2.5 ug/mL | bar: 4 to 7; 11; 13 to 15
12 | Detection titration | Show table with selected level and higher/lower levels | bar: 4 to 5; 7; 11; 13
13 | Detection Stabilizer | Biostab vs. Stabilzyme | bar: 3; 5; 7; 11; 13; 22
14 | Effect of Buffer | Test various blocking buffers for coating and analyte diluent compositions | 1: N/A; 19: N/A; 21: NA
15 | Sample dilution | Test different dilutions and impact on assay sensitivity | bar: 4 to 5; 7; 13
16 | Assay Incubation times | Test 5-5-5 against 10-10-10 | 19: N/A
17 | Whole blood spike recovery | Spike into blood with known (low) endogenous level at 8 point curve | 3: N/A; 9: N/A; 11: N/A; 12: N/A; 13: N/A; 15: N/A; 23: N/A; 26: N/A; 28: N/A; 29: N/A
18 | Hematocrit effect | Plasma from spiked blood tested, plot W.B vs. plasma | 3: N/A; 9: N/A; 11: N/A; 12: N/A; 13: N/A; 15: N/A; 23: N/A; 26: N/A; 28: N/A; 29: N/A
19 | Plasma spike recovery | Spike into plasma with known (low) endogenous level at 8 point curve | 2: N/A; 11: N/A; 12: N/A; 13: N/A; 29: N/A
20 | Stability | CAB, DAB stability, Reference tips, with a 4-point standard curve | 29: N/A
21 | Matrix effects | Spike recovery in hemolysed, lipemic, icteric samples for 6 spike levels | 21: N/A
22 | EDTA vs. Heparin |  | 29: N/A
23 | Final Clinical correlation | If Whole blood spike recovery does not show acceptable % recovery | bar: 4 to 7; 10 to 13; 27 to 28
24 | Development report completed | If Whole blood spike recovery does not show acceptable % recovery | bar: 7
25 | Release to MFG | bar: 6
Complete
in Progress
Poor response with Ab pair of current choice, re-evaluating AB screening
Need BSL2/3 lab for those steps.
Issues
N/A | Experiments not applicable to that Assay
* HbA1c: Initial work focused on Colorimetric/Enzymatic approach for assay. Issues included reagent mix formulation to match kit, high background, poor sensitivity and poor clinical correlation.
* 25 OH Vit D: Extensive optimization done for releasing agents needed in sample extraction buffer, Good response with calibrators in depleted serum however poor clinical correlation.
* Vitamin B-12: The commercial antibodies tested do not yield a dose response. Need in-house binders.
* HbSAg: Different responses seen with different subtypes of analyte, including WHO standard, Calibrators (Different subtypes & Mutant forms) need to be sent to PPD to be evaluated on Clinical analyzer before Ab selection can be finalized.
* cTnT: Re-screening to finalize antibody pair that will respond to clinicals, response seen with calibrators but no response to clinical samples.
* Lyme disease: Optimizing to use the most relevant capture surface (antigen) to detect antibodies in patient serum/plasma, issues have been poor clinical correlation with reference tests (Clinical analyzer & benhc top ELISA)

## Development & Validation plan 1
FOR ANALYTE QUANTIFICATION ASSAYS
| Experiment | Information desired | Comments
1 | MTP screen | Summary information of antibody screens, 4 point curves | Show data on all pairs, as well as summary table with color coded response
2 | Conjugation Type: SH or NH2 | Compare SH vs. NH2 conjugations (For assays with limited antibodies to screen) | MTP
3 | Cross reactivity | For 4 best Ab. Pairs:Test cross reactivity with related analytes, 6 point standard curves | On MTP, Range of the cross reactants tested will come from commercial kits or Literature based on Clinical disease and normal healthy levels
4 | Interference | For 4 best Ab. Pairs:Test interference with related analytes, 6 point standard curves | On MTP, Effect of related analyte at 3x the highest level on assay
5 | Epitope map | For 4 best Ab. Pairs | On MTP, Optional
6 | Theranos screen | For 3 best Ab. Pairs: Summary information for 6 point curves in A.B, whole blood and plasma | On readers
7 | Training set | For 1 Ab. Pair correlation of N=10 clinical samples spanning the assay range range | Confirmation that antibody pair will recognize Clinical samples
| Antibody Pair finalized | Based on the best dose response and % recovery of above experiments
8 | Plasma Screen | Screen normal plasma to see ranges | Minimum 10 samples screened
9 | Whole Blood Screen | Screen normal blood to see ranges. | Minimum 10 samples screened
10 | Capture titration | Show table with selected level and higher/lower levels | Titrate at 20, 10, 5, 2.5 ug/mL
11 | Detection titration | Show table with selected level and higher/lower levels | Titrate at 4 levels
12 | Detection Stabilizer | Biostab vs. Stabilzyme | Can be combined with Titration
13 | Effect of Buffer | Test various blocking buffers for coating and analyte diluent compositions
14 | Effect of Incubation time | Compare different assay incubation times | 2-2-2 vs. 5-5-5 vs. 10-10-10,
15 | Multiplex protocol optimization | Different dilution effects, impact of PSW
16 | % CV | Mid range analyte on 24 instruments | HT tomtec coating of 24 trays, Mean % CV and tray differential should not be >15%
17 | Precision tests | 3 lot inter and intra assay results with standard curve and 5 QC samples, 3 cartridges per lot | Standard curve + 5 QC levels at LLOQ, Low QC, Mid QC and High QC,) Need %CV on back calculated concentrations, use as many machines as available
18 | Calibrator comparison | Run our controls on reference kit and vice versa, calibrate with WHO standards | To calibrate our standards
19 | Dilution linearity | Serial dilution of high samples in buffer and sera | low endogenous plasma/ sera needed, in its absence high clinical sample can be diluted with low sample
20 | Whole blood spike recovery | Spike into blood with known (low) endogenous level at 8 point curve | show calibration of tips as well as calculated concentrations
21 | Hematocrit effect | Plasma from spiked blood tested, plot W.B vs. plasma
22 | Plasma spike recovery | Spike into plasma with known (low) endogenous level at 8 point curve | show calibration of tips as well as calculated concentrations
23 | LLOQ/ULOQ | Use 2x above ULOQ and below LLOQ to define assay limits | minimum 8-point standard curve in matrix
24 | Selectivity | Spike into W.B at 3 levels over endogenous to calculate recovery | N=10 whole blood samples, 5 male and 5 female
25 | Stability | CAB, DAB stability, Reference tips in some cases, with a 4-point standard curve | At 4ºC and RT for 24 weeks
26 | Matrix effects | Spike recovery in hemolysed or lipemic samples for 6 spike levels | % recovery back calculated from a standard curve
27 | Extended range | Test linearity and dose reposnse of standard curve at 4x of ULOQ
28 | Clinical samples | N=20, spanning the range to show correlation with kit | Calibration with 12 point curve with anchor points at 2x higher than ULOQ and 2x lower than LLOQ
29 | Development report
30 | Transfer to MFG
Validation experiment to be repeated on Mini Lab
Need Biosafety Lab for assays dealing with Infectious agents

## Development & Validation plan 2
FOR QUALITATIVE ANTIBODY ASSAYS ( Non-treponemal, HIV, Lyme, ANA, Rubella)
| Experiment | Information desired | Comments
1 | Capture surface screen | Determine optimal capture surface/ combination | Show data on all combinations, as well as summary table with color coded response
2 | Screen Detection antibodies | Determine optimal detection to use
3 | Evaluate reference kits
4 | Evaluate controls | Positive and Negative controls from different sources, WHO | On readers
5 | Training set | Test Clinicals N=20 verified on a reference kit or analyzer | Confirmation that antibody pair will recognize Clinical samples
6 | Plasma Screen | Screen normal plasma to see ranges | Minimum 10 samples screened
7 | Whole Blood Screen | Screen normal blood to see ranges. | Minimum 10 samples screened
8 | Capture titration | Optimize for capture surface formulation and concentration
9 | Detection titration | Show table with selected level and higher/lower levels | Titrate at 4 levels
10 | Detection Stabilizer | Biostab vs. Stabilzyme vs, In-house formulation where applicable | Can be combined with Titration
11 | Effect of Buffer | Test various blocking buffers for coating and analyte diluent compositions
12 | Effect of Incubation time | Compare different assay incubation times | 2-2-2 vs. 5-5-5 vs. 10-10-10
13 | Multiplex protocol optimization
14 | % CV | Mid range analyte on 24 instruments | HT tomtec coating of 24 trays, Mean % CV and tray differential should not be >15%
15 | Precision tests | 3 lot inter and intra assay results with standard curve and 5 QC samples, 3 cartridges per lot | Standard curve + 5 QC levels at LLOQ, Low QC, Mid QC and High QC,) Need %CV on back calculated concentrations, use as many machines as available
16 | Dilution linearity | Serial dilution of high samples in buffer and sera | low endogenous plasma/ sera needed, in its absence high clinical sample can be diluted with low sample
17 | LLOQ/ULOQ | Define assay ranges, sensitivity
18 | Stability | Capture surface stability, DAB stability in some cases, with a 4-point standard curve | At 4ºC and RT for 24 weeks
19 | Matrix effects | Test a variety of sample types including interfering matrices
20 | Clinical samples | N=20, spanning the assay range to show correlation with kit
21 | Development report
22 | Transfer to MFG
Validation experiment to be done on Mini Lab
Need Biosafety Lab for assays dealing with Infectious agents